## MTR_Duino 5001
1 - Record unit Serial Number and Time
- Rename Sheet based on S/N
Serial Number: | MTR_Duino 5001
Date: | 2017-07-13
2 - Paste data from calibration file here: |  |  |  | 3 - Paste data from temperature bath sheet here: |  |  | 4 - Copy first 2 columns from here to 'TempCalSummary' sheet to calculate calibration coefficients |  |  |  | 5 - After calculating calibration coefficients, paste them here |  | Statistics on Calibrated Temperature and Standard Deviation:
date/time (dd/mm/yy hh:mm:ss) | Temperature ADC Counts | Reference ADC Counts | Calibrated Temp (°C) | RefTemp | Start Time | End Time | Average Temp ADC Counts | StDev Temp ADC Counts | Average Ref ADC Counts | StDev Ref ADC Counts | 6 - Also copy calibration coefficients from here into 'Temp Calib. History' Sheet |  | RefTemp | Average Calibrated Temp (°C) | StDev Temp
2017-07-14 20:10 | 8313 | 8521 | 25.5732 | -2 | 2017-07-12 04:32 | 2017-07-12 05:32 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | AA = | 0.0012008 | -2 | #DIV/0! | #DIV/0!
2017-07-14 20:20 | 8450 | 8521 | 25.2003 | -1 | 2017-07-12 07:50 | 2017-07-12 08:50 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | BB = | 0.00052681 | -1 | #DIV/0! | #DIV/0!
2017-07-14 20:30 | 8482 | 8521 | 25.1141 | 0 | 2017-07-12 11:08 | 2017-07-12 12:08 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CC = | 1.3582e-06 | 0 | #DIV/0! | #DIV/0!
2017-07-14 20:40 | 8095 | 8521 | 26.1813 | 1 | 2017-07-12 14:27 | 2017-07-12 15:26 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 1 | #DIV/0! | #DIV/0!
2017-07-14 20:50 | 8147 | 8521 | 26.0346 | 2 | 2017-07-12 17:45 | 2017-07-12 18:45 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Maximum Residual = | -0.00035239 | 2 | #DIV/0! | #DIV/0!
2017-07-14 21:00 | 8143 | 8521 | 26.0458 | 3 | 2017-07-12 21:03 | 2017-07-12 22:03 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Standard Error = | 0.00023094 | 3 | #DIV/0! | #DIV/0!
2017-07-14 21:10 | 8221 | 8521 | 25.8276 | 4 | 2017-07-13 00:21 | 2017-07-13 01:21 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | ****After pasting, calibrated temperature will be calculated in column E |  | 4 | #DIV/0! | #DIV/0!
2017-07-14 21:20 | 8187 | 8521 | 25.9225 | 5 | 2017-07-13 03:39 | 2017-07-13 04:39 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 5 | #DIV/0! | #DIV/0!
2017-07-14 21:30 | 8064 | 8521 | 26.2693 | 6 | 2017-07-13 06:58 | 2017-07-13 07:57 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 6 | #DIV/0! | #DIV/0!
2017-07-14 21:40 | 7954 | 8521 | 26.5845 | 7 | 2017-07-13 10:16 | 2017-07-13 11:16 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 7 | #DIV/0! | #DIV/0!
2017-07-14 21:50 | 7865 | 8521 | 26.8432 | 8 | 2017-07-13 13:34 | 2017-07-13 14:34 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 8 | #DIV/0! | #DIV/0!
2017-07-14 22:00 | 7765 | 8521 | 27.1379 |  |  |  | ****Times for MTR and Temperature Bath must be synchronized for auto-fill
2017-07-14 22:10 | 7660 | 8521 | 27.4519
2017-07-14 22:20 | 7564 | 8520 | 27.7433
2017-07-14 22:30 | 7479 | 8520 | 28.0048
2017-07-14 22:40 | 7406 | 8520 | 28.2321
2017-07-14 22:50 | 7345 | 8520 | 28.424
2017-07-14 23:00 | 7269 | 8520 | 28.6656
2017-07-14 23:10 | 7189 | 8520 | 28.923
2017-07-14 23:20 | 7126 | 8520 | 29.1281
2017-07-14 23:30 | 7083 | 8520 | 29.2692
2017-07-14 23:40 | 7069 | 8520 | 29.3153
2017-07-14 23:50 | 7195 | 8520 | 28.9036
2017-07-15 | 7337 | 8520 | 28.4493
2017-07-15 00:10 | 7482 | 8520 | 27.9955
2017-07-15 00:20 | 7631 | 8520 | 27.5395
2017-07-15 00:30 | 7780 | 8521 | 27.0934
2017-07-15 00:40 | 7927 | 8521 | 26.6627
2017-07-15 00:50 | 8069 | 8521 | 26.2551
2017-07-15 01:00 | 8205 | 8521 | 25.8722
2017-07-15 01:10 | 8333 | 8521 | 25.5184
2017-07-15 01:20 | 8453 | 8521 | 25.1922
2017-07-15 01:30 | 8565 | 8521 | 24.8924
2017-07-15 01:40 | 8668 | 8521 | 24.6206
2017-07-15 01:50 | 8765 | 8521 | 24.368
2017-07-15 02:00 | 8854 | 8521 | 24.1389
2017-07-15 02:10 | 8938 | 8522 | 23.9251
2017-07-15 02:20 | 9015 | 8522 | 23.7311
2017-07-15 02:30 | 9087 | 8522 | 23.5513
2017-07-15 02:40 | 9154 | 8522 | 23.3855
2017-07-15 02:50 | 9216 | 8522 | 23.2333
2017-07-15 03:00 | 9274 | 8522 | 23.092
2017-07-15 03:10 | 9328 | 8522 | 22.9612
2017-07-15 03:20 | 9378 | 8522 | 22.841
2017-07-15 03:30 | 9424 | 8522 | 22.7309
2017-07-15 03:40 | 9466 | 8522 | 22.631
2017-07-15 03:50 | 9506 | 8522 | 22.5363
2017-07-15 04:00 | 9543 | 8522 | 22.4491
2017-07-15 04:10 | 9578 | 8522 | 22.367
2017-07-15 04:20 | 9610 | 8522 | 22.2922
2017-07-15 04:30 | 9641 | 8522 | 22.22
2017-07-15 04:40 | 9669 | 8522 | 22.155
2017-07-15 04:50 | 9695 | 8522 | 22.0949
2017-07-15 05:00 | 9719 | 8522 | 22.0395
2017-07-15 05:10 | 9742 | 8522 | 21.9866
2017-07-15 05:20 | 9764 | 8522 | 21.9361
2017-07-15 05:30 | 9786 | 8522 | 21.8858
2017-07-15 05:40 | 9809 | 8522 | 21.8333
2017-07-15 05:50 | 9833 | 8522 | 21.7786
2017-07-15 06:00 | 9856 | 8522 | 21.7264
2017-07-15 06:10 | 9880 | 8522 | 21.6721
2017-07-15 06:20 | 9904 | 8522 | 21.6179
2017-07-15 06:30 | 9927 | 8522 | 21.5661
2017-07-15 06:40 | 9951 | 8522 | 21.5122
2017-07-15 06:50 | 9974 | 8522 | 21.4606
2017-07-15 07:00 | 9997 | 8522 | 21.4093
2017-07-15 07:10 | 10020 | 8522 | 21.358
2017-07-15 07:20 | 10043 | 8522 | 21.3069
2017-07-15 07:30 | 10066 | 8522 | 21.2559
2017-07-15 07:40 | 10089 | 8522 | 21.205
2017-07-15 07:50 | 10112 | 8522 | 21.1543
2017-07-15 08:00 | 10136 | 8522 | 21.1015
2017-07-15 08:10 | 10159 | 8522 | 21.051
2017-07-15 08:20 | 10181 | 8522 | 21.0029
2017-07-15 08:30 | 10205 | 8522 | 20.9505
2017-07-15 08:40 | 10228 | 8522 | 20.9004
2017-07-15 08:50 | 10251 | 8522 | 20.8505
2017-07-15 09:00 | 10274 | 8522 | 20.8006
2017-07-15 09:10 | 10297 | 8522 | 20.7509
2017-07-15 09:20 | 10321 | 8522 | 20.6992
2017-07-15 09:30 | 10344 | 8522 | 20.6497
2017-07-15 09:40 | 10368 | 8522 | 20.5983
2017-07-15 09:50 | 10391 | 8522 | 20.5491
2017-07-15 10:00 | 10414 | 8522 | 20.5
2017-07-15 10:10 | 10438 | 8522 | 20.4489
2017-07-15 10:20 | 10462 | 8522 | 20.398
2017-07-15 10:30 | 10486 | 8522 | 20.3472
2017-07-15 10:40 | 10510 | 8522 | 20.2965
2017-07-15 10:50 | 10535 | 8522 | 20.2438
2017-07-15 11:00 | 10560 | 8522 | 20.1913
2017-07-15 11:10 | 10586 | 8522 | 20.1369
2017-07-15 11:20 | 10610 | 8522 | 20.0867
2017-07-15 11:30 | 10635 | 8522 | 20.0346
2017-07-15 11:40 | 10659 | 8522 | 19.9847
2017-07-15 11:50 | 10683 | 8522 | 19.935
2017-07-15 12:00 | 10706 | 8522 | 19.8874
2017-07-15 12:10 | 10728 | 8522 | 19.842
2017-07-15 12:20 | 10750 | 8522 | 19.7967
2017-07-15 12:30 | 10771 | 8522 | 19.7536
2017-07-15 12:40 | 10793 | 8522 | 19.7085
2017-07-15 12:50 | 10814 | 8522 | 19.6656
2017-07-15 13:00 | 10835 | 8522 | 19.6227
2017-07-15 13:10 | 10856 | 8522 | 19.58
2017-07-15 13:20 | 10877 | 8522 | 19.5373
2017-07-15 13:30 | 10898 | 8522 | 19.4948
2017-07-15 13:40 | 10918 | 8522 | 19.4543
2017-07-15 13:50 | 10937 | 8522 | 19.416
2017-07-15 14:00 | 10955 | 8522 | 19.3797
2017-07-15 14:10 | 10973 | 8522 | 19.3435
2017-07-15 14:20 | 10990 | 8522 | 19.3094
2017-07-15 14:30 | 11006 | 8522 | 19.2773
2017-07-15 14:40 | 11022 | 8522 | 19.2453
2017-07-15 14:50 | 11038 | 8522 | 19.2134
2017-07-15 15:00 | 11054 | 8522 | 19.1815
2017-07-15 15:10 | 11069 | 8522 | 19.1516
2017-07-15 15:20 | 11084 | 8522 | 19.1218
2017-07-15 15:30 | 11097 | 8523 | 19.096
2017-07-15 15:40 | 11109 | 8522 | 19.0722
2017-07-15 15:50 | 11121 | 8522 | 19.0484
2017-07-15 16:00 | 11132 | 8522 | 19.0267
2017-07-15 16:10 | 11141 | 8522 | 19.0089
2017-07-15 16:20 | 11150 | 8522 | 18.9911
2017-07-15 16:30 | 11158 | 8522 | 18.9753
2017-07-15 16:40 | 11164 | 8522 | 18.9635
2017-07-15 16:50 | 11170 | 8522 | 18.9517
2017-07-15 17:00 | 11175 | 8523 | 18.9418
2017-07-15 17:10 | 11179 | 8522 | 18.934
2017-07-15 17:20 | 11182 | 8522 | 18.9281
2017-07-15 17:30 | 11184 | 8523 | 18.9241
2017-07-15 17:40 | 11186 | 8522 | 18.9202
2017-07-15 17:50 | 11187 | 8522 | 18.9182
2017-07-15 18:00 | 11186 | 8522 | 18.9202
2017-07-15 18:10 | 11184 | 8522 | 18.9241
2017-07-15 18:20 | 11180 | 8522 | 18.932
2017-07-15 18:30 | 11176 | 8522 | 18.9399
2017-07-15 18:40 | 11170 | 8522 | 18.9517
2017-07-15 18:50 | 11164 | 8522 | 18.9635
2017-07-15 19:00 | 11157 | 8522 | 18.9773
2017-07-15 19:10 | 11149 | 8522 | 18.9931
2017-07-15 19:20 | 11141 | 8522 | 19.0089
2017-07-15 19:30 | 11131 | 8522 | 19.0286
2017-07-15 19:40 | 11122 | 8522 | 19.0464
2017-07-15 19:50 | 11112 | 8522 | 19.0662
2017-07-15 20:00 | 11102 | 8522 | 19.0861
2017-07-15 20:10 | 11092 | 8522 | 19.1059
2017-07-15 20:20 | 11082 | 8522 | 19.1257
2017-07-15 20:30 | 11072 | 8522 | 19.1456
2017-07-15 20:40 | 11062 | 8522 | 19.1655
2017-07-15 20:50 | 11052 | 8522 | 19.1854
2017-07-15 21:00 | 11042 | 8522 | 19.2054
2017-07-15 21:10 | 11031 | 8522 | 19.2273
2017-07-15 21:20 | 11019 | 8522 | 19.2513
2017-07-15 21:30 | 11005 | 8522 | 19.2793
2017-07-15 21:40 | 10991 | 8522 | 19.3074
2017-07-15 21:50 | 10977 | 8522 | 19.3355
2017-07-15 22:00 | 10961 | 8522 | 19.3676
2017-07-15 22:10 | 10945 | 8522 | 19.3998
2017-07-15 22:20 | 10928 | 8522 | 19.4341
2017-07-15 22:30 | 10912 | 8522 | 19.4665
2017-07-15 22:40 | 10894 | 8522 | 19.5029
2017-07-15 22:50 | 10876 | 8522 | 19.5394
2017-07-15 23:00 | 10856 | 8522 | 19.58
2017-07-15 23:10 | 10836 | 8522 | 19.6207
2017-07-15 23:20 | 10718 | 8522 | 19.8626
2017-07-15 23:30 | 10485 | 8522 | 20.3493
2017-07-15 23:40 | 10384 | 8522 | 20.564
2017-07-15 23:50 | 10296 | 8522 | 20.7531
2017-07-16 | 10239 | 8522 | 20.8765
2017-07-16 00:10 | 10207 | 8522 | 20.9461
2017-07-16 00:20 | 10183 | 8522 | 20.9985
2017-07-16 00:30 | 10164 | 8522 | 21.0401
2017-07-16 00:40 | 10151 | 8522 | 21.0686
2017-07-16 00:50 | 10141 | 8522 | 21.0905
2017-07-16 01:00 | 10135 | 8522 | 21.1037
2017-07-16 01:10 | 10132 | 8522 | 21.1103
2017-07-16 01:20 | 10131 | 8522 | 21.1125
2017-07-16 01:30 | 10132 | 8522 | 21.1103
2017-07-16 01:40 | 10134 | 8522 | 21.1059
2017-07-16 01:50 | 10134 | 8522 | 21.1059
2017-07-16 02:00 | 10121 | 8522 | 21.1345
2017-07-16 02:10 | 10105 | 8522 | 21.1697
2017-07-16 02:20 | 10087 | 8522 | 21.2094
2017-07-16 02:30 | 10069 | 8522 | 21.2492
2017-07-16 02:40 | 10033 | 8522 | 21.3291
2017-07-16 02:50 | 10001 | 8522 | 21.4003
2017-07-16 03:00 | 9993 | 8522 | 21.4182
2017-07-16 03:10 | 9988 | 8522 | 21.4294
2017-07-16 03:20 | 9982 | 8522 | 21.4428
2017-07-16 03:30 | 9977 | 8522 | 21.4539
2017-07-16 03:40 | 9962 | 8522 | 21.4875
2017-07-16 03:50 | 9764 | 8522 | 21.9361
2017-07-16 04:00 | 9635 | 8522 | 22.234
2017-07-16 04:10 | 9574 | 8522 | 22.3764
2017-07-16 04:20 | 9544 | 8522 | 22.4468
2017-07-16 04:30 | 9528 | 8522 | 22.4844
2017-07-16 04:40 | 9512 | 8522 | 22.5222
2017-07-16 04:50 | 9504 | 8522 | 22.5411
2017-07-16 05:00 | 9507 | 8522 | 22.534
2017-07-16 05:10 | 9518 | 8522 | 22.508
2017-07-16 05:20 | 9533 | 8522 | 22.4727
2017-07-16 05:30 | 9551 | 8522 | 22.4303
2017-07-16 05:40 | 9572 | 8522 | 22.3811
2017-07-16 05:50 | 9595 | 8522 | 22.3272
2017-07-16 06:00 | 9618 | 8522 | 22.2735
2017-07-16 06:10 | 9642 | 8522 | 22.2177
2017-07-16 06:20 | 9667 | 8522 | 22.1597
2017-07-16 06:30 | 9691 | 8522 | 22.1041
2017-07-16 06:40 | 9717 | 8522 | 22.0441
2017-07-16 06:50 | 9742 | 8522 | 21.9866
2017-07-16 07:00 | 9767 | 8522 | 21.9293
2017-07-16 07:10 | 9793 | 8522 | 21.8698
2017-07-16 07:20 | 9818 | 8522 | 21.8128
2017-07-16 07:30 | 9843 | 8522 | 21.7559
2017-07-16 07:40 | 9867 | 8522 | 21.7015
2017-07-16 07:50 | 9892 | 8522 | 21.645
2017-07-16 08:00 | 9916 | 8522 | 21.5908
2017-07-16 08:10 | 9940 | 8522 | 21.5369
2017-07-16 08:20 | 9963 | 8522 | 21.4853
2017-07-16 08:30 | 9986 | 8522 | 21.4338
2017-07-16 08:40 | 10009 | 8522 | 21.3825
2017-07-16 08:50 | 10032 | 8522 | 21.3313
2017-07-16 09:00 | 10054 | 8522 | 21.2825
2017-07-16 09:10 | 10077 | 8522 | 21.2315
2017-07-16 09:20 | 10099 | 8522 | 21.183
2017-07-16 09:30 | 10122 | 8522 | 21.1323
2017-07-16 09:40 | 10145 | 8522 | 21.0817
2017-07-16 09:50 | 10167 | 8522 | 21.0335
2017-07-16 10:00 | 10189 | 8522 | 20.9854
2017-07-16 10:10 | 10211 | 8522 | 20.9374
2017-07-16 10:20 | 10232 | 8522 | 20.8917
2017-07-16 10:30 | 10254 | 8522 | 20.8439
2017-07-16 10:40 | 10275 | 8522 | 20.7985
2017-07-16 10:50 | 10296 | 8522 | 20.7531
2017-07-16 11:00 | 10317 | 8522 | 20.7078
2017-07-16 11:10 | 10337 | 8522 | 20.6648
2017-07-16 11:20 | 10357 | 8522 | 20.6218
2017-07-16 11:30 | 10377 | 8522 | 20.579
2017-07-16 11:40 | 10396 | 8522 | 20.5384
2017-07-16 11:50 | 10415 | 8522 | 20.4979
2017-07-16 12:00 | 10434 | 8522 | 20.4574
2017-07-16 12:10 | 10453 | 8522 | 20.4171
2017-07-16 12:20 | 10472 | 8522 | 20.3768
2017-07-16 12:30 | 10491 | 8522 | 20.3366
2017-07-16 12:40 | 10509 | 8522 | 20.2986
2017-07-16 12:50 | 10527 | 8522 | 20.2607
2017-07-16 13:00 | 10545 | 8522 | 20.2228
2017-07-16 13:10 | 10563 | 8522 | 20.185
2017-07-16 13:20 | 10581 | 8522 | 20.1473
2017-07-16 13:30 | 10598 | 8522 | 20.1118
2017-07-16 13:40 | 10616 | 8522 | 20.0742
2017-07-16 13:50 | 10633 | 8522 | 20.0388
2017-07-16 14:00 | 10649 | 8522 | 20.0055
2017-07-16 14:10 | 10665 | 8522 | 19.9723
2017-07-16 14:20 | 10667 | 8522 | 19.9681
2017-07-16 14:30 | 10656 | 8522 | 19.991
2017-07-16 14:40 | 10652 | 8522 | 19.9993
2017-07-16 14:50 | 10650 | 8522 | 20.0034
2017-07-16 15:00 | 10654 | 8522 | 19.9951
2017-07-16 15:10 | 10658 | 8522 | 19.9868
2017-07-16 15:20 | 10671 | 8522 | 19.9598
2017-07-16 15:30 | 10682 | 8522 | 19.937
2017-07-16 15:40 | 10691 | 8522 | 19.9184
2017-07-16 15:50 | 10700 | 8522 | 19.8998
2017-07-16 16:00 | 10711 | 8522 | 19.8771
2017-07-16 16:10 | 10722 | 8522 | 19.8544
2017-07-16 16:20 | 10730 | 8522 | 19.8379
2017-07-16 16:30 | 10734 | 8522 | 19.8297
2017-07-16 16:40 | 10735 | 8522 | 19.8276
2017-07-16 16:50 | 10734 | 8522 | 19.8297
2017-07-16 17:00 | 10706 | 8522 | 19.8874
2017-07-16 17:10 | 10698 | 8522 | 19.9039
2017-07-16 17:20 | 10685 | 8522 | 19.9308
2017-07-16 17:30 | 10666 | 8522 | 19.9702
2017-07-16 17:40 | 10653 | 8522 | 19.9972
2017-07-16 17:50 | 10640 | 8522 | 20.0242
2017-07-16 18:00 | 10630 | 8522 | 20.045
2017-07-16 18:10 | 10618 | 8522 | 20.07
2017-07-16 18:20 | 10600 | 8522 | 20.1076
2017-07-16 18:30 | 10568 | 8522 | 20.1745
2017-07-16 18:40 | 10519 | 8522 | 20.2775
2017-07-16 18:50 | 10461 | 8522 | 20.4001
2017-07-16 19:00 | 10395 | 8522 | 20.5405
2017-07-16 19:10 | 10322 | 8522 | 20.697
2017-07-16 19:20 | 10244 | 8522 | 20.8656
2017-07-16 19:30 | 10157 | 8522 | 21.0554
2017-07-16 19:40 | 10065 | 8522 | 21.2581
2017-07-16 19:50 | 9966 | 8522 | 21.4786
2017-07-16 20:00 | 9862 | 8522 | 21.7128
2017-07-16 20:10 | 9756 | 8522 | 21.9545
2017-07-16 20:20 | 9653 | 8522 | 22.1921
2017-07-16 20:30 | 9552 | 8522 | 22.428
2017-07-16 20:40 | 9444 | 8522 | 22.6833
2017-07-16 20:50 | 9337 | 8522 | 22.9395
2017-07-16 21:00 | 9245 | 8522 | 23.1625
2017-07-16 21:10 | 9108 | 8522 | 23.4992
2017-07-16 21:20 | 9000 | 8522 | 23.7687
2017-07-16 21:30 | 8905 | 8522 | 24.0088
2017-07-16 21:40 | 8817 | 8522 | 24.2338
2017-07-16 21:50 | 8735 | 8522 | 24.4458
2017-07-16 22:00 | 8655 | 8521 | 24.6547
2017-07-16 22:10 | 8577 | 8521 | 24.8605
2017-07-16 22:20 | 8501 | 8521 | 25.0631
2017-07-16 22:30 | 8427 | 8521 | 25.2624
2017-07-16 22:40 | 8355 | 8521 | 25.4582
2017-07-16 22:50 | 8286 | 8521 | 25.6476
2017-07-16 23:00 | 8219 | 8521 | 25.8332
2017-07-16 23:10 | 8155 | 8521 | 26.0121
2017-07-16 23:20 | 8093 | 8521 | 26.187
2017-07-16 23:30 | 8033 | 8521 | 26.3576
2017-07-16 23:40 | 7977 | 8521 | 26.5182
2017-07-16 23:50 | 7922 | 8521 | 26.6772
2017-07-17 | 7870 | 8521 | 26.8286
2017-07-17 00:10 | 7821 | 8521 | 26.9723
2017-07-17 00:20 | 7776 | 8521 | 27.1053
2017-07-17 00:30 | 7732 | 8521 | 27.2361
2017-07-17 00:40 | 7692 | 8521 | 27.3557
2017-07-17 00:50 | 7654 | 8520 | 27.47
2017-07-17 01:00 | 7620 | 8520 | 27.5728
2017-07-17 01:10 | 7587 | 8520 | 27.6731
2017-07-17 01:20 | 7557 | 8520 | 27.7647
2017-07-17 01:30 | 7530 | 8520 | 27.8475
2017-07-17 01:40 | 7505 | 8520 | 27.9245
2017-07-17 01:50 | 7482 | 8520 | 27.9955
2017-07-17 02:00 | 7462 | 8520 | 28.0575
2017-07-17 02:10 | 7444 | 8520 | 28.1135
2017-07-17 02:20 | 7428 | 8520 | 28.1634
2017-07-17 02:30 | 7415 | 8520 | 28.204
2017-07-17 02:40 | 7403 | 8520 | 28.2415
2017-07-17 02:50 | 7394 | 8520 | 28.2697
2017-07-17 03:00 | 7387 | 8520 | 28.2917
2017-07-17 03:10 | 7382 | 8520 | 28.3074
2017-07-17 03:20 | 7379 | 8520 | 28.3168
2017-07-17 03:30 | 7379 | 8520 | 28.3168
2017-07-17 03:40 | 7381 | 8520 | 28.3106
2017-07-17 03:50 | 7385 | 8520 | 28.298
2017-07-17 04:00 | 7391 | 8520 | 28.2791
2017-07-17 04:10 | 7400 | 8520 | 28.2509
2017-07-17 04:20 | 7410 | 8520 | 28.2196
2017-07-17 04:30 | 7422 | 8520 | 28.1821
2017-07-17 04:40 | 7436 | 8520 | 28.1384
2017-07-17 04:50 | 7452 | 8520 | 28.0886
2017-07-17 05:00 | 7470 | 8520 | 28.0327
2017-07-17 05:10 | 7491 | 8520 | 27.9677
2017-07-17 05:20 | 7515 | 8520 | 27.8937
2017-07-17 05:30 | 7541 | 8520 | 27.8137
2017-07-17 05:40 | 7570 | 8520 | 27.725
2017-07-17 05:50 | 7603 | 8520 | 27.6244
2017-07-17 06:00 | 7639 | 8520 | 27.5153
2017-07-17 06:10 | 7679 | 8520 | 27.3947
2017-07-17 06:20 | 7722 | 8520 | 27.2659
2017-07-17 06:30 | 7767 | 8521 | 27.1319
2017-07-17 06:40 | 7816 | 8521 | 26.9871
2017-07-17 06:50 | 7868 | 8521 | 26.8345
2017-07-17 07:00 | 7922 | 8521 | 26.6772
2017-07-17 07:10 | 7979 | 8521 | 26.5125
2017-07-17 07:20 | 8038 | 8521 | 26.3433
2017-07-17 07:30 | 8100 | 8521 | 26.1672
2017-07-17 07:40 | 8165 | 8521 | 25.9841
2017-07-17 07:50 | 8231 | 8521 | 25.7998
2017-07-17 08:00 | 8299 | 8521 | 25.6117
2017-07-17 08:10 | 8369 | 8521 | 25.4199
2017-07-17 08:20 | 8441 | 8521 | 25.2245
2017-07-17 08:30 | 8514 | 8521 | 25.0283
2017-07-17 08:40 | 8588 | 8521 | 24.8314
2017-07-17 08:50 | 8664 | 8521 | 24.6311
2017-07-17 09:00 | 8741 | 8521 | 24.4302
2017-07-17 09:10 | 8819 | 8521 | 24.2287
2017-07-17 09:20 | 8897 | 8522 | 24.0292
2017-07-17 09:30 | 8977 | 8522 | 23.8266
2017-07-17 09:40 | 9057 | 8522 | 23.626
2017-07-17 09:50 | 9137 | 8522 | 23.4275
2017-07-17 10:00 | 9218 | 8522 | 23.2284
2017-07-17 10:10 | 9299 | 8522 | 23.0313
2017-07-17 10:20 | 9381 | 8522 | 22.8338
2017-07-17 10:30 | 9464 | 8522 | 22.6358
2017-07-17 10:40 | 9546 | 8522 | 22.4421
2017-07-17 10:50 | 9629 | 8522 | 22.2479
2017-07-17 11:00 | 9713 | 8522 | 22.0533
2017-07-17 11:10 | 9797 | 8522 | 21.8607
2017-07-17 11:20 | 9881 | 8522 | 21.6698
2017-07-17 11:30 | 9966 | 8522 | 21.4786
2017-07-17 11:40 | 10050 | 8522 | 21.2913
2017-07-17 11:50 | 10135 | 8522 | 21.1037
2017-07-17 12:00 | 10219 | 8522 | 20.92
2017-07-17 12:10 | 10303 | 8522 | 20.738
2017-07-17 12:20 | 10373 | 8522 | 20.5876
2017-07-17 12:30 | 10437 | 8522 | 20.4511
2017-07-17 12:40 | 10505 | 8522 | 20.307
2017-07-17 12:50 | 10574 | 8522 | 20.162
2017-07-17 13:00 | 10643 | 8522 | 20.018
2017-07-17 13:10 | 10713 | 8522 | 19.873
2017-07-17 13:20 | 10781 | 8522 | 19.7331
2017-07-17 13:30 | 10846 | 8522 | 19.6003
2017-07-17 13:40 | 10913 | 8522 | 19.4644
2017-07-17 13:50 | 10978 | 8522 | 19.3335
2017-07-17 14:00 | 11029 | 8522 | 19.2313
2017-07-17 14:10 | 11074 | 8522 | 19.1416
2017-07-17 14:20 | 11113 | 8522 | 19.0643
2017-07-17 14:30 | 11144 | 8522 | 19.003
2017-07-17 14:40 | 11165 | 8522 | 18.9615
2017-07-17 14:50 | 11172 | 8522 | 18.9478
2017-07-17 15:00 | 11166 | 8522 | 18.9596
2017-07-17 15:10 | 11155 | 8522 | 18.9812
2017-07-17 15:20 | 11146 | 8522 | 18.999
2017-07-17 15:30 | 11137 | 8522 | 19.0168
2017-07-17 15:40 | 11119 | 8522 | 19.0524
2017-07-17 15:50 | 11093 | 8522 | 19.1039
2017-07-17 16:00 | 11069 | 8522 | 19.1516
2017-07-17 16:10 | 11038 | 8522 | 19.2134
2017-07-17 16:20 | 11002 | 8522 | 19.2853
2017-07-17 16:30 | 10961 | 8522 | 19.3676
2017-07-17 16:40 | 10911 | 8522 | 19.4685
2017-07-17 16:50 | 10855 | 8522 | 19.582
2017-07-17 17:00 | 10786 | 8522 | 19.7229
2017-07-17 17:10 | 10719 | 8522 | 19.8606
2017-07-17 17:20 | 10650 | 8522 | 20.0034
2017-07-17 17:30 | 10568 | 8522 | 20.1745
2017-07-17 17:40 | 10501 | 8522 | 20.3155
2017-07-17 17:50 | 10435 | 8522 | 20.4553
2017-07-17 18:00 | 10358 | 8522 | 20.6197
2017-07-17 18:10 | 10276 | 8522 | 20.7963
2017-07-17 18:20 | 10180 | 8522 | 21.0051
2017-07-17 18:30 | 10077 | 8522 | 21.2315
2017-07-17 18:40 | 9976 | 8522 | 21.4562
2017-07-17 18:50 | 9881 | 8522 | 21.6698
2017-07-17 19:00 | 9792 | 8522 | 21.8721
2017-07-17 19:10 | 9705 | 8522 | 22.0718
2017-07-17 19:20 | 9624 | 8522 | 22.2596
2017-07-17 19:30 | 9542 | 8522 | 22.4515
2017-07-17 19:40 | 9460 | 8522 | 22.6453
2017-07-17 19:50 | 9378 | 8522 | 22.841
2017-07-17 20:00 | 9292 | 8522 | 23.0483
2017-07-17 20:10 | 9202 | 8522 | 23.2676
2017-07-17 20:20 | 9107 | 8522 | 23.5017
2017-07-17 20:30 | 9011 | 8522 | 23.7411
2017-07-17 20:40 | 8915 | 8522 | 23.9834
2017-07-17 20:50 | 8816 | 8522 | 24.2364
2017-07-17 21:00 | 8715 | 8522 | 24.4978
2017-07-17 21:10 | 8614 | 8521 | 24.7626
2017-07-17 21:20 | 8514 | 8521 | 25.0283
2017-07-17 21:30 | 8410 | 8521 | 25.3084
2017-07-17 21:40 | 8308 | 8521 | 25.587
2017-07-17 21:50 | 8210 | 8521 | 25.8583
2017-07-17 22:00 | 8114 | 8521 | 26.1276
2017-07-17 22:10 | 8023 | 8521 | 26.3862
2017-07-17 22:20 | 7938 | 8521 | 26.6308
2017-07-17 22:30 | 7860 | 8521 | 26.8579
2017-07-17 22:40 | 7789 | 8521 | 27.0668
2017-07-17 22:50 | 7723 | 8521 | 27.2629
2017-07-17 23:00 | 7661 | 8521 | 27.4489
2017-07-17 23:10 | 7605 | 8520 | 27.6183
2017-07-17 23:20 | 7554 | 8520 | 27.7739
2017-07-17 23:30 | 7507 | 8520 | 27.9183
2017-07-17 23:40 | 7467 | 8520 | 28.042
2017-07-17 23:50 | 7437 | 8520 | 28.1353
2017-07-18 | 7412 | 8520 | 28.2134
2017-07-18 00:10 | 7387 | 8520 | 28.2917
2017-07-18 00:20 | 7360 | 8520 | 28.3767
2017-07-18 00:30 | 7338 | 8520 | 28.4461
2017-07-18 00:40 | 7317 | 8520 | 28.5127
2017-07-18 00:50 | 7293 | 8520 | 28.589
2017-07-18 01:00 | 7266 | 8520 | 28.6752
2017-07-18 01:10 | 7237 | 8520 | 28.7682
2017-07-18 01:20 | 7207 | 8520 | 28.8648
2017-07-18 01:30 | 7179 | 8520 | 28.9555
2017-07-18 01:40 | 7152 | 8520 | 29.0432
2017-07-18 01:50 | 7125 | 8520 | 29.1313
2017-07-18 02:00 | 7101 | 8520 | 29.21
2017-07-18 02:10 | 7085 | 8520 | 29.2626
2017-07-18 02:20 | 7071 | 8520 | 29.3087
2017-07-18 02:30 | 7054 | 8520 | 29.3649
2017-07-18 02:40 | 7037 | 8520 | 29.4212
2017-07-18 02:50 | 7022 | 8520 | 29.471
2017-07-18 03:00 | 7010 | 8520 | 29.511
2017-07-18 03:10 | 6997 | 8520 | 29.5543
2017-07-18 03:20 | 6988 | 8520 | 29.5844
2017-07-18 03:30 | 6982 | 8520 | 29.6045
2017-07-18 03:40 | 6981 | 8520 | 29.6078
2017-07-18 03:50 | 6981 | 8520 | 29.6078
2017-07-18 04:00 | 6988 | 8520 | 29.5844
2017-07-18 04:10 | 7002 | 8520 | 29.5376
2017-07-18 04:20 | 7025 | 8520 | 29.4611
2017-07-18 04:30 | 7048 | 8520 | 29.3848
2017-07-18 04:40 | 7079 | 8520 | 29.2824
2017-07-18 04:50 | 7112 | 8520 | 29.1739
2017-07-18 05:00 | 7146 | 8520 | 29.0628
2017-07-18 05:10 | 7183 | 8520 | 28.9425
2017-07-18 05:20 | 7220 | 8520 | 28.8229
2017-07-18 05:30 | 7260 | 8520 | 28.6944
2017-07-18 05:40 | 7301 | 8520 | 28.5635
2017-07-18 05:50 | 7344 | 8520 | 28.4272
2017-07-18 06:00 | 7389 | 8520 | 28.2854
2017-07-18 06:10 | 7435 | 8520 | 28.1415
2017-07-18 06:20 | 7483 | 8520 | 27.9924
2017-07-18 06:30 | 7533 | 8520 | 27.8383
2017-07-18 06:40 | 7584 | 8520 | 27.6823
2017-07-18 06:50 | 7636 | 8520 | 27.5244
2017-07-18 07:00 | 7690 | 8521 | 27.3617
2017-07-18 07:10 | 7746 | 8521 | 27.1943
2017-07-18 07:20 | 7803 | 8521 | 27.0254
2017-07-18 07:30 | 7861 | 8521 | 26.8549
2017-07-18 07:40 | 7922 | 8521 | 26.6772
2017-07-18 07:50 | 7983 | 8521 | 26.5009
2017-07-18 08:00 | 8046 | 8521 | 26.3205
2017-07-18 08:10 | 8111 | 8521 | 26.1361
2017-07-18 08:20 | 8176 | 8521 | 25.9532
2017-07-18 08:30 | 8243 | 8521 | 25.7665
2017-07-18 08:40 | 8311 | 8521 | 25.5787
2017-07-18 08:50 | 8380 | 8521 | 25.39
2017-07-18 09:00 | 8449 | 8521 | 25.203
2017-07-18 09:10 | 8520 | 8521 | 25.0123
2017-07-18 09:20 | 8590 | 8521 | 24.8261
2017-07-18 09:30 | 8661 | 8521 | 24.639
2017-07-18 09:40 | 8732 | 8521 | 24.4535
2017-07-18 09:50 | 8803 | 8521 | 24.2698
2017-07-18 10:00 | 8874 | 8522 | 24.0878
2017-07-18 10:10 | 8945 | 8522 | 23.9074
2017-07-18 10:20 | 9017 | 8522 | 23.7261
2017-07-18 10:30 | 9090 | 8522 | 23.5439
2017-07-18 10:40 | 9163 | 8522 | 23.3634
2017-07-18 10:50 | 9236 | 8522 | 23.1845
2017-07-18 11:00 | 9309 | 8522 | 23.0071
2017-07-18 11:10 | 9383 | 8522 | 22.829
2017-07-18 11:20 | 9457 | 8522 | 22.6524
2017-07-18 11:30 | 9531 | 8522 | 22.4774
2017-07-18 11:40 | 9606 | 8522 | 22.3015
2017-07-18 11:50 | 9681 | 8522 | 22.1272
2017-07-18 12:00 | 9756 | 8522 | 21.9545
2017-07-18 12:10 | 9831 | 8522 | 21.7832
2017-07-18 12:20 | 9906 | 8522 | 21.6134
2017-07-18 12:30 | 9980 | 8522 | 21.4472
2017-07-18 12:40 | 10055 | 8522 | 21.2803
2017-07-18 12:50 | 10129 | 8522 | 21.1169
2017-07-18 13:00 | 10203 | 8522 | 20.9548
2017-07-18 13:10 | 10277 | 8522 | 20.7941
2017-07-18 13:20 | 10350 | 8522 | 20.6369
2017-07-18 13:30 | 10422 | 8522 | 20.483
2017-07-18 13:40 | 10494 | 8522 | 20.3303
2017-07-18 13:50 | 10563 | 8522 | 20.185
2017-07-18 14:00 | 10631 | 8522 | 20.0429
2017-07-18 14:10 | 10695 | 8522 | 19.9101
2017-07-18 14:20 | 10756 | 8522 | 19.7844
2017-07-18 14:30 | 10811 | 8522 | 19.6717
2017-07-18 14:40 | 10859 | 8522 | 19.5739
2017-07-18 14:50 | 10900 | 8522 | 19.4907
2017-07-18 15:00 | 10935 | 8522 | 19.42
2017-07-18 15:10 | 10965 | 8522 | 19.3596
2017-07-18 15:20 | 10991 | 8522 | 19.3074
2017-07-18 15:30 | 11010 | 8522 | 19.2693
2017-07-18 15:40 | 11023 | 8522 | 19.2433
2017-07-18 15:50 | 11029 | 8522 | 19.2313
2017-07-18 16:00 | 11029 | 8522 | 19.2313
2017-07-18 16:10 | 11021 | 8522 | 19.2473
2017-07-18 16:20 | 11005 | 8522 | 19.2793
2017-07-18 16:30 | 10982 | 8522 | 19.3254
2017-07-18 16:40 | 10951 | 8522 | 19.3878
2017-07-18 16:50 | 10912 | 8522 | 19.4665
2017-07-18 17:00 | 10865 | 8522 | 19.5617
2017-07-18 17:10 | 10811 | 8522 | 19.6717
2017-07-18 17:20 | 10750 | 8522 | 19.7967
2017-07-18 17:30 | 10682 | 8522 | 19.937
2017-07-18 17:40 | 10608 | 8522 | 20.0909
2017-07-18 17:50 | 10529 | 8522 | 20.2565
2017-07-18 18:00 | 10445 | 8522 | 20.4341
2017-07-18 18:10 | 10357 | 8522 | 20.6218
2017-07-18 18:20 | 10265 | 8522 | 20.8201
2017-07-18 18:30 | 10170 | 8522 | 21.0269
2017-07-18 18:40 | 10075 | 8522 | 21.236
2017-07-18 18:50 | 9981 | 8522 | 21.445
2017-07-18 19:00 | 9889 | 8522 | 21.6517
2017-07-18 19:10 | 9801 | 8522 | 21.8515
2017-07-18 19:20 | 9713 | 8522 | 22.0533
2017-07-18 19:30 | 9624 | 8522 | 22.2596
2017-07-18 19:40 | 9533 | 8522 | 22.4727
2017-07-18 19:50 | 9438 | 8522 | 22.6976
2017-07-18 20:00 | 9341 | 8522 | 22.9299
2017-07-18 20:10 | 9240 | 8522 | 23.1747
2017-07-18 20:20 | 9136 | 8522 | 23.4299
2017-07-18 20:30 | 9029 | 8522 | 23.696
2017-07-18 20:40 | 8919 | 8522 | 23.9732
2017-07-18 20:50 | 8806 | 8522 | 24.2621
2017-07-18 21:00 | 8691 | 8521 | 24.5604
2017-07-18 21:10 | 8574 | 8521 | 24.8685
2017-07-18 21:20 | 8456 | 8521 | 25.1841
2017-07-18 21:30 | 8337 | 8521 | 25.5074
2017-07-18 21:40 | 8220 | 8521 | 25.8304
2017-07-18 21:50 | 8106 | 8521 | 26.1502
2017-07-18 22:00 | 7996 | 8521 | 26.4636
2017-07-18 22:10 | 7892 | 8521 | 26.7644
2017-07-18 22:20 | 7795 | 8521 | 27.049
2017-07-18 22:30 | 7705 | 8521 | 27.3167
2017-07-18 22:40 | 7623 | 8521 | 27.5637
2017-07-18 22:50 | 7546 | 8520 | 27.7984
2017-07-18 23:00 | 7477 | 8520 | 28.011
2017-07-18 23:10 | 7413 | 8520 | 28.2102
2017-07-18 23:20 | 7355 | 8520 | 28.3924
2017-07-18 23:30 | 7303 | 8520 | 28.5572
2017-07-18 23:40 | 7256 | 8520 | 28.7072
2017-07-18 23:50 | 7213 | 8520 | 28.8455
2017-07-19 | 7175 | 8520 | 28.9684
2017-07-19 00:10 | 7141 | 8520 | 29.0791
2017-07-19 00:20 | 7111 | 8520 | 29.1772
2017-07-19 00:30 | 7084 | 8520 | 29.2659
2017-07-19 00:40 | 7060 | 8520 | 29.3451
2017-07-19 00:50 | 7039 | 8520 | 29.4146
2017-07-19 01:00 | 7021 | 8520 | 29.4744
2017-07-19 01:10 | 7006 | 8520 | 29.5243
2017-07-19 01:20 | 6992 | 8520 | 29.571
2017-07-19 01:30 | 6981 | 8520 | 29.6078
2017-07-19 01:40 | 6973 | 8520 | 29.6346
2017-07-19 01:50 | 6966 | 8520 | 29.6581
2017-07-19 02:00 | 6962 | 8520 | 29.6715
2017-07-19 02:10 | 6959 | 8520 | 29.6816
2017-07-19 02:20 | 6959 | 8520 | 29.6816
2017-07-19 02:30 | 6960 | 8520 | 29.6782
2017-07-19 02:40 | 6963 | 8520 | 29.6682
2017-07-19 02:50 | 6969 | 8520 | 29.648
2017-07-19 03:00 | 6976 | 8520 | 29.6246
2017-07-19 03:10 | 6985 | 8520 | 29.5944
2017-07-19 03:20 | 6997 | 8520 | 29.5543
2017-07-19 03:30 | 7025 | 8520 | 29.4611
2017-07-19 03:40 | 7143 | 8520 | 29.0725
2017-07-19 03:50 | 7249 | 8520 | 28.7297
2017-07-19 04:00 | 7354 | 8520 | 28.3956
2017-07-19 04:10 | 7457 | 8520 | 28.0731
2017-07-19 04:20 | 7560 | 8520 | 27.7555
2017-07-19 04:30 | 7662 | 8520 | 27.4459
2017-07-19 04:40 | 7763 | 8521 | 27.1438
2017-07-19 04:50 | 7864 | 8521 | 26.8462
2017-07-19 05:00 | 7962 | 8521 | 26.5614
2017-07-19 05:10 | 8059 | 8521 | 26.2835
2017-07-19 05:20 | 8153 | 8521 | 26.0177
2017-07-19 05:30 | 8245 | 8521 | 25.761
2017-07-19 05:40 | 8337 | 8521 | 25.5074
2017-07-19 05:50 | 8426 | 8521 | 25.2651
2017-07-19 06:00 | 8513 | 8521 | 25.031
2017-07-19 06:10 | 8598 | 8521 | 24.8049
2017-07-19 06:20 | 8681 | 8521 | 24.5866
2017-07-19 06:30 | 8762 | 8521 | 24.3757
2017-07-19 06:40 | 8840 | 8521 | 24.1748
2017-07-19 06:50 | 8918 | 8522 | 23.9758
2017-07-19 07:00 | 8994 | 8522 | 23.7838
2017-07-19 07:10 | 9069 | 8522 | 23.5961
2017-07-19 07:20 | 9143 | 8522 | 23.4127
2017-07-19 07:30 | 9216 | 8522 | 23.2333
2017-07-19 07:40 | 9290 | 8522 | 23.0531
2017-07-19 07:50 | 9366 | 8522 | 22.8698
2017-07-19 08:00 | 9441 | 8522 | 22.6904
2017-07-19 08:10 | 9516 | 8522 | 22.5127
2017-07-19 08:20 | 9589 | 8522 | 22.3413
2017-07-19 08:30 | 9661 | 8522 | 22.1736
2017-07-19 08:40 | 9731 | 8522 | 22.0119
2017-07-19 08:50 | 9799 | 8522 | 21.8561
2017-07-19 09:00 | 9868 | 8522 | 21.6992
2017-07-19 09:10 | 9937 | 8522 | 21.5436
2017-07-19 09:20 | 10006 | 8522 | 21.3892
2017-07-19 09:30 | 10074 | 8522 | 21.2382
2017-07-19 09:40 | 10141 | 8522 | 21.0905
2017-07-19 09:50 | 10206 | 8522 | 20.9483
2017-07-19 10:00 | 10271 | 8522 | 20.8071
2017-07-19 10:10 | 10335 | 8522 | 20.6691
2017-07-19 10:20 | 10398 | 8522 | 20.5341
2017-07-19 10:30 | 10460 | 8522 | 20.4022
2017-07-19 10:40 | 10520 | 8522 | 20.2754
2017-07-19 10:50 | 10579 | 8522 | 20.1515
2017-07-19 11:00 | 10638 | 8522 | 20.0284
2017-07-19 11:10 | 10699 | 8522 | 19.9019
2017-07-19 11:20 | 10760 | 8522 | 19.7762
2017-07-19 11:30 | 10821 | 8522 | 19.6513
2017-07-19 11:40 | 10881 | 8522 | 19.5292
2017-07-19 11:50 | 10941 | 8522 | 19.4079
2017-07-19 12:00 | 11001 | 8522 | 19.2873
2017-07-19 12:10 | 11077 | 8522 | 19.1357
2017-07-19 12:20 | 11171 | 8522 | 18.9497
2017-07-19 12:30 | 11221 | 8522 | 18.8515
2017-07-19 12:40 | 11271 | 8522 | 18.7538
2017-07-19 12:50 | 11324 | 8522 | 18.6508
2017-07-19 13:00 | 11378 | 8522 | 18.5463
2017-07-19 13:10 | 11429 | 8522 | 18.4482
2017-07-19 13:20 | 11477 | 8522 | 18.3563
2017-07-19 13:30 | 11499 | 8522 | 18.3144
2017-07-19 13:40 | 11431 | 8522 | 18.4444
2017-07-19 13:50 | 11271 | 8522 | 18.7538
2017-07-19 14:00 | 11042 | 8522 | 19.2054
2017-07-19 14:10 | 10803 | 8522 | 19.6881
2017-07-19 14:20 | 10605 | 8522 | 20.0971
2017-07-19 14:30 | 10464 | 8522 | 20.3937
2017-07-19 14:40 | 10374 | 8522 | 20.5854
2017-07-19 14:50 | 10355 | 8522 | 20.6261
2017-07-19 15:00 | 10289 | 8522 | 20.7682
2017-07-19 15:10 | 10159 | 8522 | 21.051
2017-07-19 15:20 | 10040 | 8522 | 21.3135
2017-07-19 15:30 | 9966 | 8522 | 21.4786
2017-07-19 15:40 | 9921 | 8522 | 21.5796
2017-07-19 15:50 | 9875 | 8522 | 21.6834
2017-07-19 16:00 | 9785 | 8522 | 21.8881
2017-07-19 16:10 | 9643 | 8522 | 22.2154
2017-07-19 16:20 | 9471 | 8522 | 22.6192
2017-07-19 16:30 | 9309 | 8522 | 23.0071
2017-07-19 16:40 | 9261 | 8522 | 23.1236
2017-07-19 16:50 | 9239 | 8522 | 23.1771
2017-07-19 17:00 | 9212 | 8522 | 23.2431
2017-07-19 17:10 | 9137 | 8522 | 23.4275
2017-07-19 17:20 | 9266 | 8522 | 23.1114
2017-07-19 17:30 | 9413 | 8522 | 22.7572
2017-07-19 17:40 | 9476 | 8522 | 22.6073
2017-07-19 17:50 | 9494 | 8522 | 22.5647
2017-07-19 18:00 | 9513 | 8522 | 22.5198
2017-07-19 18:10 | 9527 | 8522 | 22.4868
2017-07-19 18:20 | 9540 | 8522 | 22.4562
2017-07-19 18:30 | 9569 | 8522 | 22.3881
2017-07-19 18:40 | 9669 | 8522 | 22.155
2017-07-19 18:50 | 9891 | 8522 | 21.6472
2017-07-19 19:00 | 10213 | 8522 | 20.9331
2017-07-19 19:10 | 10601 | 8522 | 20.1055
2017-07-19 19:20 | 11036 | 8522 | 19.2174
2017-07-19 19:30 | 11508 | 8522 | 18.2972
2017-07-19 19:40 | 12001 | 8522 | 17.3805
2017-07-19 19:50 | 12508 | 8522 | 16.4809
2017-07-19 20:00 | 13026 | 8522 | 15.6032
2017-07-19 20:10 | 13556 | 8522 | 14.7449
2017-07-19 20:20 | 14089 | 8522 | 13.919
2017-07-19 20:30 | 14621 | 8522 | 13.1289
2017-07-19 20:40 | 15149 | 8522 | 12.3761
2017-07-19 20:50 | 15668 | 8522 | 11.6644
2017-07-19 21:00 | 16124 | 8522 | 11.0605
2017-07-19 21:10 | 16509 | 8522 | 10.5654
2017-07-19 21:20 | 16691 | 8522 | 10.3359
2017-07-19 21:30 | 16236 | 8522 | 10.9151
2017-07-19 21:40 | 15510 | 8522 | 11.8782
2017-07-19 21:50 | 15080 | 8522 | 12.4728
2017-07-19 22:00 | 14749 | 8522 | 12.9437
2017-07-19 22:10 | 14524 | 8522 | 13.2705
2017-07-19 22:20 | 14325 | 8522 | 13.5644
2017-07-19 22:30 | 14139 | 8522 | 13.8433
2017-07-19 22:40 | 13978 | 8522 | 14.088
2017-07-19 22:50 | 13826 | 8522 | 14.322
2017-07-19 23:00 | 13614 | 8522 | 14.6532
2017-07-19 23:10 | 13400 | 8522 | 14.9935
2017-07-19 23:20 | 13202 | 8523 | 15.3139
2017-07-19 23:30 | 13019 | 8522 | 15.6148
2017-07-19 23:40 | 12860 | 8522 | 15.8801
2017-07-19 23:50 | 12716 | 8522 | 16.1236
2017-07-20 | 12581 | 8523 | 16.3547
2017-07-20 00:10 | 12453 | 8523 | 16.5764
2017-07-20 00:20 | 12331 | 8522 | 16.7902
2017-07-20 00:30 | 12226 | 8523 | 16.9761
2017-07-20 00:40 | 12125 | 8523 | 17.1565
2017-07-20 00:50 | 12038 | 8523 | 17.3134
2017-07-20 01:00 | 11953 | 8523 | 17.4678
2017-07-20 01:10 | 11871 | 8522 | 17.618
2017-07-20 01:20 | 11795 | 8522 | 17.7583
2017-07-20 01:30 | 11723 | 8522 | 17.8921
2017-07-20 01:40 | 11654 | 8522 | 18.0212
2017-07-20 01:50 | 11620 | 8522 | 18.0851
2017-07-20 02:00 | 11586 | 8522 | 18.1493
2017-07-20 02:10 | 11670 | 8522 | 17.9912
2017-07-20 02:20 | 11583 | 8522 | 18.1549
2017-07-20 02:30 | 11529 | 8522 | 18.2573
2017-07-20 02:40 | 11477 | 8522 | 18.3563
2017-07-20 02:50 | 11429 | 8522 | 18.4482
2017-07-20 03:00 | 11375 | 8522 | 18.5521
2017-07-20 03:10 | 11323 | 8522 | 18.6527
2017-07-20 03:20 | 11274 | 8522 | 18.748
2017-07-20 03:30 | 11226 | 8522 | 18.8417
2017-07-20 03:40 | 11182 | 8522 | 18.9281
2017-07-20 03:50 | 11140 | 8522 | 19.0109
2017-07-20 04:00 | 11100 | 8522 | 19.09
2017-07-20 04:10 | 11062 | 8522 | 19.1655
2017-07-20 04:20 | 10960 | 8522 | 19.3697
2017-07-20 04:30 | 10845 | 8522 | 19.6024
2017-07-20 04:40 | 10779 | 8522 | 19.7372
2017-07-20 04:50 | 10720 | 8522 | 19.8585
2017-07-20 05:00 | 10658 | 8522 | 19.9868
2017-07-20 05:10 | 10519 | 8522 | 20.2775
2017-07-20 05:20 | 10438 | 8522 | 20.4489
2017-07-20 05:30 | 10406 | 8522 | 20.5171
2017-07-20 05:40 | 10388 | 8522 | 20.5555
2017-07-20 05:50 | 10375 | 8522 | 20.5833
2017-07-20 06:00 | 10365 | 8522 | 20.6047
2017-07-20 06:10 | 10359 | 8522 | 20.6176
2017-07-20 06:20 | 10356 | 8522 | 20.624
2017-07-20 06:30 | 10355 | 8522 | 20.6261
2017-07-20 06:40 | 10352 | 8522 | 20.6326
2017-07-20 06:50 | 10350 | 8522 | 20.6369
2017-07-20 07:00 | 10347 | 8522 | 20.6433
2017-07-20 07:10 | 10342 | 8522 | 20.654
2017-07-20 07:20 | 10337 | 8522 | 20.6648
2017-07-20 07:30 | 10331 | 8522 | 20.6777
2017-07-20 07:40 | 10325 | 8522 | 20.6906
2017-07-20 07:50 | 10319 | 8522 | 20.7035
2017-07-20 08:00 | 10313 | 8522 | 20.7164
2017-07-20 08:10 | 10307 | 8522 | 20.7293
2017-07-20 08:20 | 10300 | 8522 | 20.7444
2017-07-20 08:30 | 10294 | 8522 | 20.7574
2017-07-20 08:40 | 10288 | 8522 | 20.7704
2017-07-20 08:50 | 10282 | 8522 | 20.7833
2017-07-20 09:00 | 10276 | 8522 | 20.7963
2017-07-20 09:10 | 10270 | 8522 | 20.8093
2017-07-20 09:20 | 10264 | 8522 | 20.8223
2017-07-20 09:30 | 10259 | 8522 | 20.8331
2017-07-20 09:40 | 10253 | 8522 | 20.8461
2017-07-20 09:50 | 10248 | 8522 | 20.857
2017-07-20 10:00 | 10242 | 8522 | 20.87
2017-07-20 10:10 | 10237 | 8522 | 20.8808
2017-07-20 10:20 | 10232 | 8522 | 20.8917
2017-07-20 10:30 | 10227 | 8522 | 20.9026
2017-07-20 10:40 | 10222 | 8522 | 20.9135
2017-07-20 10:50 | 10217 | 8522 | 20.9243
2017-07-20 11:00 | 10212 | 8522 | 20.9352
2017-07-20 11:10 | 10207 | 8522 | 20.9461
2017-07-20 11:20 | 10202 | 8522 | 20.957
2017-07-20 11:30 | 10197 | 8522 | 20.9679
2017-07-20 11:40 | 10193 | 8522 | 20.9767
2017-07-20 11:50 | 10188 | 8522 | 20.9876
2017-07-20 12:00 | 10184 | 8522 | 20.9963
2017-07-20 12:10 | 10180 | 8522 | 21.0051
2017-07-20 12:20 | 10176 | 8522 | 21.0138
2017-07-20 12:30 | 10173 | 8522 | 21.0204
2017-07-20 12:40 | 10169 | 8522 | 21.0291
2017-07-20 12:50 | 10166 | 8522 | 21.0357
2017-07-20 13:00 | 10164 | 8522 | 21.0401
2017-07-20 13:10 | 10161 | 8522 | 21.0467
2017-07-20 13:20 | 10158 | 8522 | 21.0532
2017-07-20 13:30 | 10156 | 8522 | 21.0576
2017-07-20 13:40 | 10154 | 8522 | 21.062
2017-07-20 13:50 | 10151 | 8522 | 21.0686
2017-07-20 14:00 | 10149 | 8522 | 21.073
2017-07-20 14:10 | 10145 | 8522 | 21.0817
2017-07-20 14:20 | 10140 | 8522 | 21.0927
2017-07-20 14:30 | 10134 | 8522 | 21.1059
2017-07-20 14:40 | 10127 | 8522 | 21.1213
2017-07-20 14:50 | 10119 | 8522 | 21.1389
2017-07-20 15:00 | 10130 | 8522 | 21.1147
2017-07-20 15:10 | 10637 | 8522 | 20.0305
2017-07-20 15:20 | 11065 | 8522 | 19.1596
2017-07-20 15:30 | 11537 | 8522 | 18.2421
2017-07-20 15:40 | 11661 | 8522 | 18.008
2017-07-20 15:50 | 11531 | 8522 | 18.2535
2017-07-20 16:00 | 11417 | 8522 | 18.4713
2017-07-20 16:10 | 11311 | 8522 | 18.676
2017-07-20 16:20 | 11232 | 8522 | 18.83
2017-07-20 16:30 | 11144 | 8522 | 19.003
2017-07-20 16:40 | 10878 | 8522 | 19.5353
2017-07-20 16:50 | 10783 | 8522 | 19.729
2017-07-20 17:00 | 10689 | 8522 | 19.9226
2017-07-20 17:10 | 10646 | 8522 | 20.0117
2017-07-20 17:20 | 10637 | 8522 | 20.0305
2017-07-20 17:30 | 10637 | 8522 | 20.0305
2017-07-20 17:40 | 10639 | 8522 | 20.0263
2017-07-20 17:50 | 10642 | 8522 | 20.0201
2017-07-20 18:00 | 10644 | 8522 | 20.0159
2017-07-20 18:10 | 10647 | 8522 | 20.0097
2017-07-20 18:20 | 10648 | 8522 | 20.0076
2017-07-20 18:30 | 10649 | 8522 | 20.0055
2017-07-20 18:40 | 10650 | 8522 | 20.0034
2017-07-20 18:50 | 10651 | 8522 | 20.0013
2017-07-20 19:00 | 10652 | 8522 | 19.9993
2017-07-20 19:10 | 10652 | 8522 | 19.9993
2017-07-20 19:20 | 10653 | 8522 | 19.9972
2017-07-20 19:30 | 10653 | 8522 | 19.9972
2017-07-20 19:40 | 10654 | 8522 | 19.9951
2017-07-20 19:50 | 10652 | 8522 | 19.9993
2017-07-20 20:00 | 10652 | 8522 | 19.9993
2017-07-20 20:10 | 10653 | 8522 | 19.9972
2017-07-20 20:20 | 10652 | 8522 | 19.9993
2017-07-20 20:30 | 10652 | 8522 | 19.9993
2017-07-20 20:40 | 10648 | 8522 | 20.0076
2017-07-20 20:50 | 10645 | 8522 | 20.0138
2017-07-20 21:00 | 10646 | 8522 | 20.0117
2017-07-20 21:10 | 10644 | 8522 | 20.0159
2017-07-20 21:20 | 10644 | 8522 | 20.0159
2017-07-20 21:30 | 10641 | 8522 | 20.0221
2017-07-20 21:40 | 10723 | 8522 | 19.8523
2017-07-20 21:50 | 10764 | 8522 | 19.768
2017-07-20 22:00 | 10781 | 8522 | 19.7331
2017-07-20 22:10 | 10789 | 8522 | 19.7167
2017-07-20 22:20 | 10801 | 8522 | 19.6922
2017-07-20 22:30 | 10811 | 8522 | 19.6717
2017-07-20 22:40 | 10806 | 8522 | 19.6819
2017-07-20 22:50 | 10797 | 8522 | 19.7003
2017-07-20 23:00 | 10782 | 8522 | 19.731
2017-07-20 23:10 | 10758 | 8522 | 19.7803
2017-07-20 23:20 | 10739 | 8522 | 19.8194
2017-07-20 23:30 | 10721 | 8522 | 19.8564
2017-07-20 23:40 | 10701 | 8522 | 19.8977
2017-07-20 23:50 | 10679 | 8522 | 19.9433
2017-07-21 | 10656 | 8522 | 19.991
2017-07-21 00:10 | 10632 | 8522 | 20.0409
2017-07-21 00:20 | 10555 | 8522 | 20.2018
2017-07-21 00:30 | 10300 | 8522 | 20.7444
2017-07-21 00:40 | 10117 | 8522 | 21.1433
2017-07-21 00:50 | 10096 | 8522 | 21.1896
2017-07-21 01:00 | 10095 | 8522 | 21.1918
2017-07-21 01:10 | 10103 | 8522 | 21.1741
2017-07-21 01:20 | 10106 | 8522 | 21.1675
2017-07-21 01:30 | 10097 | 8522 | 21.1874
2017-07-21 01:40 | 10082 | 8522 | 21.2205
2017-07-21 01:50 | 10060 | 8522 | 21.2692
2017-07-21 02:00 | 10047 | 8522 | 21.298
2017-07-21 02:10 | 10035 | 8522 | 21.3246
2017-07-21 02:20 | 10025 | 8522 | 21.3469
2017-07-21 02:30 | 10014 | 8522 | 21.3714
2017-07-21 02:40 | 10004 | 8522 | 21.3936
2017-07-21 02:50 | 9994 | 8522 | 21.416
2017-07-21 03:00 | 9982 | 8522 | 21.4428
2017-07-21 03:10 | 9970 | 8522 | 21.4696
2017-07-21 03:20 | 9963 | 8522 | 21.4853
2017-07-21 03:30 | 9959 | 8522 | 21.4942
2017-07-21 03:40 | 9956 | 8522 | 21.501
2017-07-21 03:50 | 9959 | 8522 | 21.4942
2017-07-21 04:00 | 9962 | 8522 | 21.4875
2017-07-21 04:10 | 9963 | 8522 | 21.4853
2017-07-21 04:20 | 9953 | 8522 | 21.5077
2017-07-21 04:30 | 9942 | 8522 | 21.5324
2017-07-21 04:40 | 9933 | 8522 | 21.5526
2017-07-21 04:50 | 9913 | 8522 | 21.5976
2017-07-21 05:00 | 9858 | 8522 | 21.7219
2017-07-21 05:10 | 9807 | 8522 | 21.8378
2017-07-21 05:20 | 9812 | 8522 | 21.8264
2017-07-21 05:30 | 9853 | 8522 | 21.7332
2017-07-21 05:40 | 9895 | 8522 | 21.6382
2017-07-21 05:50 | 9938 | 8522 | 21.5414
2017-07-21 06:00 | 9982 | 8522 | 21.4428
2017-07-21 06:10 | 10028 | 8522 | 21.3402
2017-07-21 06:20 | 10076 | 8522 | 21.2338
2017-07-21 06:30 | 10125 | 8522 | 21.1257
2017-07-21 06:40 | 10169 | 8522 | 21.0291
2017-07-21 06:50 | 10204 | 8522 | 20.9527
2017-07-21 07:00 | 10240 | 8522 | 20.8743
2017-07-21 07:10 | 10268 | 8522 | 20.8136
2017-07-21 07:20 | 10301 | 8522 | 20.7423
2017-07-21 07:30 | 10338 | 8522 | 20.6626
2017-07-21 07:40 | 10373 | 8522 | 20.5876
2017-07-21 07:50 | 10406 | 8522 | 20.5171
2017-07-21 08:00 | 10436 | 8522 | 20.4532
2017-07-21 08:10 | 10467 | 8522 | 20.3874
2017-07-21 08:20 | 10496 | 8522 | 20.326
2017-07-21 08:30 | 10528 | 8522 | 20.2586
2017-07-21 08:40 | 10559 | 8522 | 20.1934
2017-07-21 08:50 | 10579 | 8522 | 20.1515
2017-07-21 09:00 | 10611 | 8522 | 20.0846
2017-07-21 09:10 | 10645 | 8522 | 20.0138
2017-07-21 09:20 | 10681 | 8522 | 19.9391
2017-07-21 09:30 | 10718 | 8522 | 19.8626
2017-07-21 09:40 | 10751 | 8522 | 19.7947
2017-07-21 09:50 | 10783 | 8522 | 19.729
2017-07-21 10:00 | 10813 | 8522 | 19.6676
2017-07-21 10:10 | 10841 | 8522 | 19.6105
2017-07-21 10:20 | 10869 | 8522 | 19.5536
2017-07-21 10:30 | 10899 | 8522 | 19.4928
2017-07-21 10:40 | 10929 | 8522 | 19.4321
2017-07-21 10:50 | 10960 | 8522 | 19.3697
2017-07-21 11:00 | 10987 | 8522 | 19.3154
2017-07-21 11:10 | 11014 | 8522 | 19.2613
2017-07-21 11:20 | 11040 | 8522 | 19.2094
2017-07-21 11:30 | 11067 | 8522 | 19.1556
2017-07-21 11:40 | 11091 | 8522 | 19.1079
2017-07-21 11:50 | 11117 | 8522 | 19.0563
2017-07-21 12:00 | 11142 | 8522 | 19.0069
2017-07-21 12:10 | 11167 | 8522 | 18.9576
2017-07-21 12:20 | 11190 | 8522 | 18.9123
2017-07-21 12:30 | 11214 | 8522 | 18.8652
2017-07-21 12:40 | 11238 | 8522 | 18.8182
2017-07-21 12:50 | 11262 | 8522 | 18.7714
2017-07-21 13:00 | 11286 | 8522 | 18.7246
2017-07-21 13:10 | 11311 | 8522 | 18.676
2017-07-21 13:20 | 11336 | 8522 | 18.6275
2017-07-21 13:30 | 11361 | 8522 | 18.5792
2017-07-21 13:40 | 11380 | 8522 | 18.5425
2017-07-21 13:50 | 11397 | 8522 | 18.5097
2017-07-21 14:00 | 11404 | 8522 | 18.4963
2017-07-21 14:10 | 11415 | 8522 | 18.4751
2017-07-21 14:20 | 11421 | 8522 | 18.4636
2017-07-21 14:30 | 11400 | 8522 | 18.504
2017-07-21 14:40 | 11402 | 8522 | 18.5001
2017-07-21 14:50 | 11407 | 8522 | 18.4905
2017-07-21 15:00 | 11397 | 8522 | 18.5097
2017-07-21 15:10 | 11278 | 8522 | 18.7402
2017-07-21 15:20 | 11261 | 8522 | 18.7733
2017-07-21 15:30 | 11255 | 8522 | 18.785
2017-07-21 15:40 | 11247 | 8522 | 18.8007
2017-07-21 15:50 | 11245 | 8522 | 18.8046
2017-07-21 16:00 | 11246 | 8522 | 18.8026
2017-07-21 16:10 | 11248 | 8522 | 18.7987
2017-07-21 16:20 | 11253 | 8522 | 18.7889
2017-07-21 16:30 | 11259 | 8522 | 18.7772
2017-07-21 16:40 | 11266 | 8522 | 18.7636
2017-07-21 16:50 | 11273 | 8522 | 18.7499
2017-07-21 17:00 | 11280 | 8522 | 18.7363
2017-07-21 17:10 | 11284 | 8522 | 18.7285
2017-07-21 17:20 | 11287 | 8522 | 18.7226
2017-07-21 17:30 | 11289 | 8522 | 18.7188
2017-07-21 17:40 | 11289 | 8522 | 18.7188
2017-07-21 17:50 | 11292 | 8522 | 18.7129
2017-07-21 18:00 | 11296 | 8522 | 18.7051
2017-07-21 18:10 | 11301 | 8522 | 18.6954
2017-07-21 18:20 | 11307 | 8522 | 18.6838
2017-07-21 18:30 | 11311 | 8522 | 18.676
2017-07-21 18:40 | 11314 | 8522 | 18.6702
2017-07-21 18:50 | 11318 | 8522 | 18.6624
2017-07-21 19:00 | 11322 | 8522 | 18.6547
2017-07-21 19:10 | 11327 | 8522 | 18.645
2017-07-21 19:20 | 11332 | 8522 | 18.6353
2017-07-21 19:30 | 11337 | 8522 | 18.6256
2017-07-21 19:40 | 11341 | 8522 | 18.6178
2017-07-21 19:50 | 11343 | 8522 | 18.614
2017-07-21 20:00 | 11342 | 8522 | 18.6159
2017-07-21 20:10 | 11332 | 8522 | 18.6353
2017-07-21 20:20 | 11320 | 8522 | 18.6585
2017-07-21 20:30 | 11310 | 8522 | 18.6779
2017-07-21 20:40 | 11269 | 8522 | 18.7577
2017-07-21 20:50 | 11241 | 8522 | 18.8124
2017-07-21 21:00 | 11190 | 8522 | 18.9123
2017-07-21 21:10 | 11136 | 8522 | 19.0188
2017-07-21 21:20 | 11086 | 8522 | 19.1178
2017-07-21 21:30 | 11038 | 8522 | 19.2134
2017-07-21 21:40 | 10990 | 8522 | 19.3094
2017-07-21 21:50 | 10944 | 8522 | 19.4019
2017-07-21 22:00 | 10900 | 8522 | 19.4907
2017-07-21 22:10 | 10858 | 8522 | 19.5759
2017-07-21 22:20 | 10820 | 8522 | 19.6533
2017-07-21 22:30 | 10785 | 8522 | 19.7249
2017-07-21 22:40 | 10754 | 8522 | 19.7885
2017-07-21 22:50 | 10729 | 8522 | 19.84
2017-07-21 23:00 | 10707 | 8522 | 19.8853
2017-07-21 23:10 | 10689 | 8522 | 19.9226
2017-07-21 23:20 | 10675 | 8522 | 19.9515
2017-07-21 23:30 | 10665 | 8522 | 19.9723
2017-07-21 23:40 | 10661 | 8522 | 19.9806
2017-07-21 23:50 | 10660 | 8522 | 19.9827
2017-07-22 | 10663 | 8522 | 19.9764
2017-07-22 00:10 | 10668 | 8522 | 19.9661
2017-07-22 00:20 | 10665 | 8522 | 19.9723
2017-07-22 00:30 | 10657 | 8522 | 19.9889
2017-07-22 00:40 | 10650 | 8522 | 20.0034
2017-07-22 00:50 | 10646 | 8522 | 20.0117
2017-07-22 01:00 | 10645 | 8522 | 20.0138
2017-07-22 01:10 | 10647 | 8522 | 20.0097
2017-07-22 01:20 | 10650 | 8522 | 20.0034
2017-07-22 01:30 | 10656 | 8522 | 19.991
2017-07-22 01:40 | 10665 | 8522 | 19.9723
2017-07-22 01:50 | 10675 | 8522 | 19.9515
2017-07-22 02:00 | 10688 | 8522 | 19.9246
2017-07-22 02:10 | 10680 | 8522 | 19.9412
2017-07-22 02:20 | 10662 | 8522 | 19.9785
2017-07-22 02:30 | 10663 | 8522 | 19.9764
2017-07-22 02:40 | 10662 | 8522 | 19.9785
2017-07-22 02:50 | 10664 | 8522 | 19.9744
2017-07-22 03:00 | 10666 | 8522 | 19.9702
2017-07-22 03:10 | 10664 | 8522 | 19.9744
2017-07-22 03:20 | 10652 | 8522 | 19.9993
2017-07-22 03:30 | 10631 | 8522 | 20.0429
2017-07-22 03:40 | 10601 | 8522 | 20.1055
2017-07-22 03:50 | 10569 | 8522 | 20.1724
2017-07-22 04:00 | 10537 | 8522 | 20.2396
2017-07-22 04:10 | 10505 | 8522 | 20.307
2017-07-22 04:20 | 10474 | 8522 | 20.3726
2017-07-22 04:30 | 10447 | 8522 | 20.4298
2017-07-22 04:40 | 10426 | 8522 | 20.4745
2017-07-22 04:50 | 10403 | 8522 | 20.5235
2017-07-22 05:00 | 10376 | 8522 | 20.5811
2017-07-22 05:10 | 10347 | 8522 | 20.6433
2017-07-22 05:20 | 10338 | 8522 | 20.6626
2017-07-22 05:30 | 10336 | 8522 | 20.6669
2017-07-22 05:40 | 10340 | 8522 | 20.6583
2017-07-22 05:50 | 10348 | 8522 | 20.6412
2017-07-22 06:00 | 10359 | 8522 | 20.6176
2017-07-22 06:10 | 10374 | 8522 | 20.5854
2017-07-22 06:20 | 10392 | 8522 | 20.5469
2017-07-22 06:30 | 10412 | 8522 | 20.5043
2017-07-22 06:40 | 10435 | 8522 | 20.4553
2017-07-22 06:50 | 10459 | 8522 | 20.4043
2017-07-22 07:00 | 10485 | 8522 | 20.3493
2017-07-22 07:10 | 10512 | 8522 | 20.2923
2017-07-22 07:20 | 10539 | 8522 | 20.2354
2017-07-22 07:30 | 10567 | 8522 | 20.1766
2017-07-22 07:40 | 10596 | 8522 | 20.1159
2017-07-22 07:50 | 10626 | 8522 | 20.0534
2017-07-22 08:00 | 10654 | 8522 | 19.9951
2017-07-22 08:10 | 10684 | 8522 | 19.9329
2017-07-22 08:20 | 10714 | 8522 | 19.8709
2017-07-22 08:30 | 10742 | 8522 | 19.8132
2017-07-22 08:40 | 10769 | 8522 | 19.7577
2017-07-22 08:50 | 10795 | 8522 | 19.7044
2017-07-22 09:00 | 10819 | 8522 | 19.6554
2017-07-22 09:10 | 10843 | 8522 | 19.6064
2017-07-22 09:20 | 10868 | 8522 | 19.5556
2017-07-22 09:30 | 10892 | 8522 | 19.5069
2017-07-22 09:40 | 10917 | 8522 | 19.4563
2017-07-22 09:50 | 10941 | 8522 | 19.4079
2017-07-22 10:00 | 10966 | 8522 | 19.3576
2017-07-22 10:10 | 10993 | 8522 | 19.3034
2017-07-22 10:20 | 11020 | 8522 | 19.2493
2017-07-22 10:30 | 11047 | 8522 | 19.1954
2017-07-22 10:40 | 11073 | 8522 | 19.1436
2017-07-22 10:50 | 11098 | 8522 | 19.094
2017-07-22 11:00 | 11123 | 8522 | 19.0445
2017-07-22 11:10 | 11148 | 8522 | 18.9951
2017-07-22 11:20 | 11172 | 8522 | 18.9478
2017-07-22 11:30 | 11196 | 8522 | 18.9006
2017-07-22 11:40 | 11220 | 8522 | 18.8535
2017-07-22 11:50 | 11245 | 8522 | 18.8046
2017-07-22 12:00 | 11269 | 8522 | 18.7577
2017-07-22 12:10 | 11292 | 8522 | 18.7129
2017-07-22 12:20 | 11315 | 8522 | 18.6682
2017-07-22 12:30 | 11338 | 8522 | 18.6236
2017-07-22 12:40 | 11361 | 8522 | 18.5792
2017-07-22 12:50 | 11386 | 8522 | 18.5309
2017-07-22 13:00 | 11412 | 8522 | 18.4809
2017-07-22 13:10 | 11440 | 8522 | 18.4271
2017-07-22 13:20 | 11469 | 8522 | 18.3716
2017-07-22 13:30 | 11499 | 8522 | 18.3144
2017-07-22 13:40 | 11529 | 8522 | 18.2573
2017-07-22 13:50 | 11559 | 8522 | 18.2004
2017-07-22 14:00 | 11589 | 8522 | 18.1436
2017-07-22 14:10 | 11618 | 8522 | 18.0889
2017-07-22 14:20 | 11647 | 8522 | 18.0343
2017-07-22 14:30 | 11664 | 8522 | 18.0024
2017-07-22 14:40 | 11668 | 8522 | 17.9949
2017-07-22 14:50 | 11679 | 8522 | 17.9743
2017-07-22 15:00 | 11692 | 8522 | 17.95
2017-07-22 15:10 | 11705 | 8522 | 17.9257
2017-07-22 15:20 | 11716 | 8522 | 17.9051
2017-07-22 15:30 | 11726 | 8522 | 17.8865
2017-07-22 15:40 | 11732 | 8522 | 17.8753
2017-07-22 15:50 | 11723 | 8522 | 17.8921
2017-07-22 16:00 | 11709 | 8522 | 17.9182
2017-07-22 16:10 | 11702 | 8522 | 17.9313
2017-07-22 16:20 | 11701 | 8522 | 17.9332
2017-07-22 16:30 | 11699 | 8522 | 17.9369
2017-07-22 16:40 | 11655 | 8522 | 18.0193
2017-07-22 16:50 | 11631 | 8522 | 18.0644
2017-07-22 17:00 | 11607 | 8522 | 18.1096
2017-07-22 17:10 | 11580 | 8522 | 18.1606
2017-07-22 17:20 | 11543 | 8522 | 18.2307
2017-07-22 17:30 | 11496 | 8522 | 18.3201
2017-07-22 17:40 | 11447 | 8522 | 18.4137
2017-07-22 17:50 | 11399 | 8522 | 18.5059
2017-07-22 18:00 | 11341 | 8522 | 18.6178
2017-07-22 18:10 | 11291 | 8522 | 18.7149
2017-07-22 18:20 | 11248 | 8522 | 18.7987
2017-07-22 18:30 | 11207 | 8522 | 18.879
2017-07-22 18:40 | 11167 | 8522 | 18.9576
2017-07-22 18:50 | 11123 | 8522 | 19.0445
2017-07-22 19:00 | 11066 | 8522 | 19.1576
2017-07-22 19:10 | 11018 | 8522 | 19.2533
2017-07-22 19:20 | 10970 | 8522 | 19.3495
2017-07-22 19:30 | 10925 | 8522 | 19.4402
2017-07-22 19:40 | 10887 | 8522 | 19.5171
2017-07-22 19:50 | 10851 | 8522 | 19.5902
2017-07-22 20:00 | 10815 | 8522 | 19.6635
2017-07-22 20:10 | 10779 | 8522 | 19.7372
2017-07-22 20:20 | 10740 | 8522 | 19.8173
2017-07-22 20:30 | 10705 | 8522 | 19.8895
2017-07-22 20:40 | 10675 | 8522 | 19.9515
2017-07-22 20:50 | 10649 | 8522 | 20.0055
2017-07-22 21:00 | 10628 | 8522 | 20.0492
2017-07-22 21:10 | 10609 | 8522 | 20.0888
2017-07-22 21:20 | 10592 | 8522 | 20.1243
2017-07-22 21:30 | 10571 | 8522 | 20.1683
2017-07-22 21:40 | 10551 | 8522 | 20.2102
2017-07-22 21:50 | 10528 | 8522 | 20.2586
2017-07-22 22:00 | 10508 | 8522 | 20.3007
2017-07-22 22:10 | 10490 | 8522 | 20.3387
2017-07-22 22:20 | 10474 | 8522 | 20.3726
2017-07-22 22:30 | 10461 | 8522 | 20.4001
2017-07-22 22:40 | 10451 | 8522 | 20.4213
2017-07-22 22:50 | 10439 | 8522 | 20.4468
2017-07-22 23:00 | 10428 | 8522 | 20.4702
2017-07-22 23:10 | 10419 | 8522 | 20.4894
2017-07-22 23:20 | 10411 | 8522 | 20.5064
2017-07-22 23:30 | 10405 | 8522 | 20.5192
2017-07-22 23:40 | 10399 | 8522 | 20.532
2017-07-22 23:50 | 10395 | 8522 | 20.5405
2017-07-23 | 10388 | 8522 | 20.5555
2017-07-23 00:10 | 10375 | 8522 | 20.5833
2017-07-23 00:20 | 10358 | 8522 | 20.6197
2017-07-23 00:30 | 10342 | 8522 | 20.654
2017-07-23 00:40 | 10330 | 8522 | 20.6798
2017-07-23 00:50 | 10324 | 8522 | 20.6927
2017-07-23 01:00 | 10322 | 8522 | 20.697
2017-07-23 01:10 | 10324 | 8522 | 20.6927
2017-07-23 01:20 | 10332 | 8522 | 20.6755
2017-07-23 01:30 | 10344 | 8522 | 20.6497
2017-07-23 01:40 | 10360 | 8522 | 20.6154
2017-07-23 01:50 | 10371 | 8522 | 20.5919
2017-07-23 02:00 | 10356 | 8522 | 20.624
2017-07-23 02:10 | 10346 | 8522 | 20.6454
2017-07-23 02:20 | 10348 | 8522 | 20.6412
2017-07-23 02:30 | 10352 | 8522 | 20.6326
2017-07-23 02:40 | 10357 | 8522 | 20.6218
2017-07-23 02:50 | 10363 | 8522 | 20.609
2017-07-23 03:00 | 10367 | 8522 | 20.6004
2017-07-23 03:10 | 10372 | 8522 | 20.5897
2017-07-23 03:20 | 10385 | 8522 | 20.5619
2017-07-23 03:30 | 10394 | 8522 | 20.5427
2017-07-23 03:40 | 10397 | 8522 | 20.5363
2017-07-23 03:50 | 10401 | 8522 | 20.5277
2017-07-23 04:00 | 10414 | 8522 | 20.5
2017-07-23 04:10 | 10434 | 8522 | 20.4574
2017-07-23 04:20 | 10456 | 8522 | 20.4107
2017-07-23 04:30 | 10478 | 8522 | 20.3641
2017-07-23 04:40 | 10498 | 8522 | 20.3218
2017-07-23 04:50 | 10517 | 8522 | 20.2817
2017-07-23 05:00 | 10534 | 8522 | 20.2459
2017-07-23 05:10 | 10549 | 8522 | 20.2144
2017-07-23 05:20 | 10560 | 8522 | 20.1913
2017-07-23 05:30 | 10569 | 8522 | 20.1724
2017-07-23 05:40 | 10560 | 8522 | 20.1913
2017-07-23 05:50 | 10535 | 8522 | 20.2438
2017-07-23 06:00 | 10513 | 8522 | 20.2902
2017-07-23 06:10 | 10494 | 8522 | 20.3303
2017-07-23 06:20 | 10486 | 8522 | 20.3472
2017-07-23 06:30 | 10489 | 8522 | 20.3408
2017-07-23 06:40 | 10500 | 8522 | 20.3176
2017-07-23 06:50 | 10518 | 8522 | 20.2796
2017-07-23 07:00 | 10540 | 8522 | 20.2333
2017-07-23 07:10 | 10565 | 8522 | 20.1808
2017-07-23 07:20 | 10592 | 8522 | 20.1243
2017-07-23 07:30 | 10620 | 8522 | 20.0659
2017-07-23 07:40 | 10649 | 8522 | 20.0055
2017-07-23 07:50 | 10678 | 8522 | 19.9453
2017-07-23 08:00 | 10707 | 8522 | 19.8853
2017-07-23 08:10 | 10735 | 8522 | 19.8276
2017-07-23 08:20 | 10763 | 8522 | 19.77
2017-07-23 08:30 | 10790 | 8522 | 19.7147
2017-07-23 08:40 | 10817 | 8522 | 19.6595
2017-07-23 08:50 | 10843 | 8522 | 19.6064
2017-07-23 09:00 | 10869 | 8522 | 19.5536
2017-07-23 09:10 | 10894 | 8522 | 19.5029
2017-07-23 09:20 | 10919 | 8522 | 19.4523
2017-07-23 09:30 | 10943 | 8522 | 19.4039
2017-07-23 09:40 | 10966 | 8522 | 19.3576
2017-07-23 09:50 | 10989 | 8522 | 19.3114
2017-07-23 10:00 | 11011 | 8522 | 19.2673
2017-07-23 10:10 | 11034 | 8522 | 19.2213
2017-07-23 10:20 | 11057 | 8522 | 19.1755
2017-07-23 10:30 | 11080 | 8522 | 19.1297
2017-07-23 10:40 | 11102 | 8522 | 19.0861
2017-07-23 10:50 | 11124 | 8522 | 19.0425
2017-07-23 11:00 | 11145 | 8522 | 19.001
2017-07-23 11:10 | 11165 | 8522 | 18.9615
2017-07-23 11:20 | 11185 | 8522 | 18.9222
2017-07-23 11:30 | 11203 | 8522 | 18.8868
2017-07-23 11:40 | 11220 | 8522 | 18.8535
2017-07-23 11:50 | 11236 | 8522 | 18.8222
2017-07-23 12:00 | 11250 | 8522 | 18.7948
2017-07-23 12:10 | 11262 | 8522 | 18.7714
2017-07-23 12:20 | 11273 | 8522 | 18.7499
2017-07-23 12:30 | 11284 | 8522 | 18.7285
2017-07-23 12:40 | 11294 | 8522 | 18.709
2017-07-23 12:50 | 11303 | 8522 | 18.6915
2017-07-23 13:00 | 11312 | 8522 | 18.6741
2017-07-23 13:10 | 11322 | 8522 | 18.6547
2017-07-23 13:20 | 11333 | 8522 | 18.6333
2017-07-23 13:30 | 11345 | 8522 | 18.6101
2017-07-23 13:40 | 11357 | 8522 | 18.5869
2017-07-23 13:50 | 11369 | 8522 | 18.5637
2017-07-23 14:00 | 11381 | 8522 | 18.5406
2017-07-23 14:10 | 11394 | 8522 | 18.5155
2017-07-23 14:20 | 11408 | 8522 | 18.4886
2017-07-23 14:30 | 11423 | 8522 | 18.4597
2017-07-23 14:40 | 11438 | 8522 | 18.431
2017-07-23 14:50 | 11453 | 8522 | 18.4022
2017-07-23 15:00 | 11465 | 8522 | 18.3793
2017-07-23 15:10 | 11472 | 8522 | 18.3659
2017-07-23 15:20 | 11476 | 8522 | 18.3582
2017-07-23 15:30 | 11479 | 8522 | 18.3525
2017-07-23 15:40 | 11483 | 8522 | 18.3449
2017-07-23 15:50 | 11488 | 8522 | 18.3353
2017-07-23 16:00 | 11493 | 8522 | 18.3258
2017-07-23 16:10 | 11498 | 8522 | 18.3163
2017-07-23 16:20 | 11500 | 8522 | 18.3125
2017-07-23 16:30 | 11497 | 8522 | 18.3182
2017-07-23 16:40 | 11489 | 8522 | 18.3334
2017-07-23 16:50 | 11473 | 8522 | 18.364
2017-07-23 17:00 | 11454 | 8522 | 18.4003
2017-07-23 17:10 | 11430 | 8522 | 18.4463
2017-07-23 17:20 | 11408 | 8522 | 18.4886
2017-07-23 17:30 | 11391 | 8522 | 18.5213
2017-07-23 17:40 | 11374 | 8522 | 18.5541
2017-07-23 17:50 | 11356 | 8522 | 18.5888
2017-07-23 18:00 | 11315 | 8522 | 18.6682
2017-07-23 18:10 | 11269 | 8522 | 18.7577
2017-07-23 18:20 | 11223 | 8522 | 18.8476
2017-07-23 18:30 | 11169 | 8522 | 18.9537
2017-07-23 18:40 | 11120 | 8522 | 19.0504
2017-07-23 18:50 | 11068 | 8522 | 19.1536
2017-07-23 19:00 | 11009 | 8522 | 19.2713
2017-07-23 19:10 | 10957 | 8522 | 19.3757
2017-07-23 19:20 | 10909 | 8522 | 19.4725
2017-07-23 19:30 | 10864 | 8522 | 19.5637
2017-07-23 19:40 | 10822 | 8522 | 19.6493
2017-07-23 19:50 | 10785 | 8522 | 19.7249
2017-07-23 20:00 | 10751 | 8522 | 19.7947
2017-07-23 20:10 | 10716 | 8522 | 19.8668
2017-07-23 20:20 | 10678 | 8522 | 19.9453
2017-07-23 20:30 | 10639 | 8522 | 20.0263
2017-07-23 20:40 | 10599 | 8522 | 20.1097
2017-07-23 20:50 | 10548 | 8522 | 20.2165
2017-07-23 21:00 | 10481 | 8522 | 20.3577
2017-07-23 21:10 | 10430 | 8522 | 20.4659
2017-07-23 21:20 | 10399 | 8522 | 20.532
2017-07-23 21:30 | 10378 | 8522 | 20.5769
2017-07-23 21:40 | 10381 | 8522 | 20.5705
2017-07-23 21:50 | 10394 | 8522 | 20.5427
2017-07-23 22:00 | 10408 | 8522 | 20.5128
2017-07-23 22:10 | 10418 | 8522 | 20.4915
2017-07-23 22:20 | 10402 | 8522 | 20.5256
2017-07-23 22:30 | 10395 | 8522 | 20.5405
2017-07-23 22:40 | 10409 | 8522 | 20.5107
2017-07-23 22:50 | 10434 | 8522 | 20.4574
2017-07-23 23:00 | 10463 | 8522 | 20.3959
2017-07-23 23:10 | 10491 | 8522 | 20.3366
2017-07-23 23:20 | 10517 | 8522 | 20.2817
2017-07-23 23:30 | 10541 | 8522 | 20.2312
2017-07-23 23:40 | 10563 | 8522 | 20.185
2017-07-23 23:50 | 10569 | 8522 | 20.1724
2017-07-24 | 10560 | 8522 | 20.1913
2017-07-24 00:10 | 10537 | 8522 | 20.2396
2017-07-24 00:20 | 10529 | 8522 | 20.2565
2017-07-24 00:30 | 10538 | 8522 | 20.2375
2017-07-24 00:40 | 10555 | 8522 | 20.2018
2017-07-24 00:50 | 10574 | 8522 | 20.162
2017-07-24 01:00 | 10592 | 8522 | 20.1243
2017-07-24 01:10 | 10586 | 8522 | 20.1369
2017-07-24 01:20 | 10573 | 8522 | 20.1641
2017-07-24 01:30 | 10563 | 8522 | 20.185
2017-07-24 01:40 | 10559 | 8522 | 20.1934
2017-07-24 01:50 | 10567 | 8522 | 20.1766
2017-07-24 02:00 | 10598 | 8522 | 20.1118
2017-07-24 02:10 | 10640 | 8522 | 20.0242
2017-07-24 02:20 | 10684 | 8522 | 19.9329
2017-07-24 02:30 | 10728 | 8522 | 19.842
2017-07-24 02:40 | 10769 | 8522 | 19.7577
2017-07-24 02:50 | 10791 | 8522 | 19.7126
2017-07-24 03:00 | 10806 | 8522 | 19.6819
2017-07-24 03:10 | 10811 | 8522 | 19.6717
2017-07-24 03:20 | 10815 | 8522 | 19.6635
2017-07-24 03:30 | 10826 | 8522 | 19.6411
2017-07-24 03:40 | 10856 | 8522 | 19.58
2017-07-24 03:50 | 10897 | 8522 | 19.4968
2017-07-24 04:00 | 10933 | 8522 | 19.424
2017-07-24 04:10 | 10955 | 8522 | 19.3797
2017-07-24 04:20 | 10983 | 8522 | 19.3234
2017-07-24 04:30 | 11020 | 8522 | 19.2493
2017-07-24 04:40 | 11066 | 8522 | 19.1576
2017-07-24 04:50 | 11117 | 8522 | 19.0563
2017-07-24 05:00 | 11169 | 8522 | 18.9537
2017-07-24 05:10 | 11217 | 8522 | 18.8594
2017-07-24 05:20 | 11256 | 8522 | 18.7831
2017-07-24 05:30 | 11289 | 8522 | 18.7188
2017-07-24 05:40 | 11323 | 8522 | 18.6527
2017-07-24 05:50 | 11361 | 8522 | 18.5792
2017-07-24 06:00 | 11401 | 8522 | 18.502
2017-07-24 06:10 | 11442 | 8522 | 18.4233
2017-07-24 06:20 | 11481 | 8522 | 18.3487
2017-07-24 06:30 | 11521 | 8522 | 18.2725
2017-07-24 06:40 | 11559 | 8522 | 18.2004
2017-07-24 06:50 | 11598 | 8522 | 18.1266
2017-07-24 07:00 | 11637 | 8522 | 18.0531
2017-07-24 07:10 | 11677 | 8522 | 17.9781
2017-07-24 07:20 | 11718 | 8522 | 17.9014
2017-07-24 07:30 | 11760 | 8522 | 17.8232
2017-07-24 07:40 | 11802 | 8522 | 17.7453
2017-07-24 07:50 | 11844 | 8522 | 17.6677
2017-07-24 08:00 | 11885 | 8522 | 17.5923
2017-07-24 08:10 | 11926 | 8522 | 17.5172
2017-07-24 08:20 | 11966 | 8522 | 17.4441
2017-07-24 08:30 | 12005 | 8522 | 17.3732
2017-07-24 08:40 | 12043 | 8522 | 17.3043
2017-07-24 08:50 | 12080 | 8522 | 17.2375
2017-07-24 09:00 | 12116 | 8522 | 17.1727
2017-07-24 09:10 | 12151 | 8522 | 17.1099
2017-07-24 09:20 | 12187 | 8522 | 17.0455
2017-07-24 09:30 | 12221 | 8522 | 16.9849
2017-07-24 09:40 | 12254 | 8522 | 16.9263
2017-07-24 09:50 | 12286 | 8522 | 16.8696
2017-07-24 10:00 | 12317 | 8523 | 16.8149
2017-07-24 10:10 | 12348 | 8522 | 16.7603
2017-07-24 10:20 | 12377 | 8522 | 16.7093
2017-07-24 10:30 | 12405 | 8522 | 16.6603
2017-07-24 10:40 | 12433 | 8522 | 16.6113
2017-07-24 10:50 | 12462 | 8522 | 16.5608
2017-07-24 11:00 | 12490 | 8522 | 16.5121
2017-07-24 11:10 | 12517 | 8522 | 16.4653
2017-07-24 11:20 | 12544 | 8522 | 16.4186
2017-07-24 11:30 | 12570 | 8523 | 16.3737
2017-07-24 11:40 | 12596 | 8522 | 16.3289
2017-07-24 11:50 | 12620 | 8522 | 16.2877
2017-07-24 12:00 | 12644 | 8522 | 16.2465
2017-07-24 12:10 | 12667 | 8522 | 16.2072
2017-07-24 12:20 | 12689 | 8522 | 16.1696
2017-07-24 12:30 | 12711 | 8523 | 16.1321
2017-07-24 12:40 | 12733 | 8522 | 16.0947
2017-07-24 12:50 | 12754 | 8522 | 16.059
2017-07-24 13:00 | 12773 | 8522 | 16.0268
2017-07-24 13:10 | 12791 | 8522 | 15.9964
2017-07-24 13:20 | 12809 | 8522 | 15.966
2017-07-24 13:30 | 12827 | 8522 | 15.9356
2017-07-24 13:40 | 12844 | 8523 | 15.907
2017-07-24 13:50 | 12857 | 8523 | 15.8851
2017-07-24 14:00 | 12867 | 8523 | 15.8683
2017-07-24 14:10 | 12875 | 8522 | 15.8549
2017-07-24 14:20 | 12879 | 8522 | 15.8482
2017-07-24 14:30 | 12882 | 8523 | 15.8432
2017-07-24 14:40 | 12884 | 8522 | 15.8398
2017-07-24 14:50 | 12884 | 8522 | 15.8398
2017-07-24 15:00 | 12883 | 8522 | 15.8415
2017-07-24 15:10 | 12852 | 8523 | 15.8935
2017-07-24 15:20 | 12820 | 8523 | 15.9474
2017-07-24 15:30 | 12800 | 8523 | 15.9812
2017-07-24 15:40 | 12767 | 8522 | 16.037
2017-07-24 15:50 | 12730 | 8522 | 16.0998
2017-07-24 16:00 | 12693 | 8522 | 16.1628
2017-07-24 16:10 | 12654 | 8522 | 16.2294
2017-07-24 16:20 | 12618 | 8523 | 16.2911
2017-07-24 16:30 | 12583 | 8522 | 16.3513
2017-07-24 16:40 | 12540 | 8522 | 16.4255
2017-07-24 16:50 | 12509 | 8522 | 16.4791
2017-07-24 17:00 | 12482 | 8522 | 16.526
2017-07-24 17:10 | 12455 | 8522 | 16.573
2017-07-24 17:20 | 12428 | 8522 | 16.6201
2017-07-24 17:30 | 12397 | 8522 | 16.6743
2017-07-24 17:40 | 12364 | 8523 | 16.7321
2017-07-24 17:50 | 12336 | 8522 | 16.7814
2017-07-24 18:00 | 12312 | 8522 | 16.8237
2017-07-24 18:10 | 12292 | 8522 | 16.859
2017-07-24 18:20 | 12274 | 8522 | 16.8909
2017-07-24 18:30 | 12259 | 8522 | 16.9174
2017-07-24 18:40 | 12246 | 8522 | 16.9405
2017-07-24 18:50 | 12193 | 8522 | 17.0348
2017-07-24 19:00 | 12117 | 8522 | 17.1709
2017-07-24 19:10 | 12077 | 8522 | 17.2429
2017-07-24 19:20 | 12057 | 8522 | 17.279
2017-07-24 19:30 | 12045 | 8522 | 17.3007
2017-07-24 19:40 | 12037 | 8522 | 17.3152
2017-07-24 19:50 | 12030 | 8522 | 17.3279
2017-07-24 20:00 | 12003 | 8522 | 17.3768
2017-07-24 20:10 | 11966 | 8522 | 17.4441
2017-07-24 20:20 | 11928 | 8522 | 17.5135
2017-07-24 20:30 | 11894 | 8522 | 17.5758
2017-07-24 20:40 | 11855 | 8522 | 17.6475
2017-07-24 20:50 | 11823 | 8523 | 17.7065
2017-07-24 21:00 | 11816 | 8522 | 17.7194
2017-07-24 21:10 | 11822 | 8522 | 17.7083
2017-07-24 21:20 | 11831 | 8522 | 17.6917
2017-07-24 21:30 | 11838 | 8522 | 17.6788
2017-07-24 21:40 | 11820 | 8522 | 17.712
2017-07-24 21:50 | 11791 | 8522 | 17.7657
2017-07-24 22:00 | 11777 | 8522 | 17.7916
2017-07-24 22:10 | 11771 | 8522 | 17.8028
2017-07-24 22:20 | 11768 | 8522 | 17.8083
2017-07-24 22:30 | 11768 | 8522 | 17.8083
2017-07-24 22:40 | 11771 | 8522 | 17.8028
2017-07-24 22:50 | 11771 | 8522 | 17.8028
2017-07-24 23:00 | 11767 | 8522 | 17.8102
2017-07-24 23:10 | 11774 | 8522 | 17.7972
2017-07-24 23:20 | 11784 | 8522 | 17.7787
2017-07-24 23:30 | 11779 | 8522 | 17.7879
2017-07-24 23:40 | 11778 | 8522 | 17.7898
2017-07-24 23:50 | 11761 | 8522 | 17.8213
2017-07-25 | 11734 | 8522 | 17.8716
2017-07-25 00:10 | 11694 | 8522 | 17.9462
2017-07-25 00:20 | 11611 | 8522 | 18.1021
2017-07-25 00:30 | 11544 | 8522 | 18.2288
2017-07-25 00:40 | 11531 | 8522 | 18.2535
2017-07-25 00:50 | 11543 | 8522 | 18.2307
2017-07-25 01:00 | 11508 | 8522 | 18.2972
2017-07-25 01:10 | 11408 | 8522 | 18.4886
2017-07-25 01:20 | 11297 | 8522 | 18.7032
2017-07-25 01:30 | 11206 | 8522 | 18.8809
2017-07-25 01:40 | 11134 | 8522 | 19.0227
2017-07-25 01:50 | 11110 | 8522 | 19.0702
2017-07-25 02:00 | 11152 | 8522 | 18.9872
2017-07-25 02:10 | 11210 | 8522 | 18.8731
2017-07-25 02:20 | 11165 | 8522 | 18.9615
2017-07-25 02:30 | 11097 | 8522 | 19.096
2017-07-25 02:40 | 11051 | 8522 | 19.1874
2017-07-25 02:50 | 11044 | 8522 | 19.2014
2017-07-25 03:00 | 11004 | 8522 | 19.2813
2017-07-25 03:10 | 10954 | 8522 | 19.3817
2017-07-25 03:20 | 10934 | 8522 | 19.422
2017-07-25 03:30 | 10936 | 8522 | 19.418
2017-07-25 03:40 | 10981 | 8522 | 19.3275
2017-07-25 03:50 | 11027 | 8522 | 19.2353
2017-07-25 04:00 | 10987 | 8522 | 19.3154
2017-07-25 04:10 | 10921 | 8522 | 19.4483
2017-07-25 04:20 | 10854 | 8522 | 19.5841
2017-07-25 04:30 | 10804 | 8522 | 19.686
2017-07-25 04:40 | 10770 | 8522 | 19.7557
2017-07-25 04:50 | 10725 | 8522 | 19.8482
2017-07-25 05:00 | 10719 | 8522 | 19.8606
2017-07-25 05:10 | 10767 | 8522 | 19.7618
2017-07-25 05:20 | 10810 | 8522 | 19.6738
2017-07-25 05:30 | 10879 | 8522 | 19.5333
2017-07-25 05:40 | 10941 | 8522 | 19.4079
2017-07-25 05:50 | 11000 | 8522 | 19.2893
2017-07-25 06:00 | 11059 | 8522 | 19.1715
2017-07-25 06:10 | 11115 | 8522 | 19.0603
2017-07-25 06:20 | 11171 | 8522 | 18.9497
2017-07-25 06:30 | 11226 | 8522 | 18.8417
2017-07-25 06:40 | 11281 | 8522 | 18.7343
2017-07-25 06:50 | 11335 | 8522 | 18.6295
2017-07-25 07:00 | 11387 | 8522 | 18.529
2017-07-25 07:10 | 11439 | 8522 | 18.4291
2017-07-25 07:20 | 11490 | 8522 | 18.3315
2017-07-25 07:30 | 11539 | 8522 | 18.2383
2017-07-25 07:40 | 11588 | 8522 | 18.1455
2017-07-25 07:50 | 11635 | 8522 | 18.0569
2017-07-25 08:00 | 11682 | 8522 | 17.9687
2017-07-25 08:10 | 11727 | 8522 | 17.8846
2017-07-25 08:20 | 11769 | 8522 | 17.8065
2017-07-25 08:30 | 11810 | 8522 | 17.7305
2017-07-25 08:40 | 11850 | 8522 | 17.6567
2017-07-25 08:50 | 11889 | 8522 | 17.585
2017-07-25 09:00 | 11928 | 8522 | 17.5135
2017-07-25 09:10 | 11966 | 8522 | 17.4441
2017-07-25 09:20 | 12005 | 8522 | 17.3732
2017-07-25 09:30 | 12042 | 8523 | 17.3061
2017-07-25 09:40 | 12077 | 8522 | 17.2429
2017-07-25 09:50 | 12112 | 8522 | 17.1799
2017-07-25 10:00 | 12146 | 8522 | 17.1189
2017-07-25 10:10 | 12180 | 8523 | 17.058
2017-07-25 10:20 | 12214 | 8522 | 16.9974
2017-07-25 10:30 | 12248 | 8522 | 16.937
2017-07-25 10:40 | 12281 | 8522 | 16.8785
2017-07-25 10:50 | 12314 | 8523 | 16.8202
2017-07-25 11:00 | 12348 | 8522 | 16.7603
2017-07-25 11:10 | 12381 | 8523 | 16.7023
2017-07-25 11:20 | 12413 | 8522 | 16.6463
2017-07-25 11:30 | 12445 | 8523 | 16.5904
2017-07-25 11:40 | 12477 | 8522 | 16.5347
2017-07-25 11:50 | 12509 | 8522 | 16.4791
2017-07-25 12:00 | 12540 | 8522 | 16.4255
2017-07-25 12:10 | 12571 | 8522 | 16.372
2017-07-25 12:20 | 12601 | 8522 | 16.3203
2017-07-25 12:30 | 12631 | 8523 | 16.2688
2017-07-25 12:40 | 12659 | 8522 | 16.2208
2017-07-25 12:50 | 12687 | 8522 | 16.173
2017-07-25 13:00 | 12714 | 8522 | 16.127
2017-07-25 13:10 | 12740 | 8523 | 16.0828
2017-07-25 13:20 | 12765 | 8522 | 16.0404
2017-07-25 13:30 | 12788 | 8522 | 16.0015
2017-07-25 13:40 | 12809 | 8522 | 15.966
2017-07-25 13:50 | 12829 | 8523 | 15.9323
2017-07-25 14:00 | 12846 | 8522 | 15.9036
2017-07-25 14:10 | 12861 | 8522 | 15.8784
2017-07-25 14:20 | 12874 | 8523 | 15.8566
2017-07-25 14:30 | 12886 | 8523 | 15.8365
2017-07-25 14:40 | 12898 | 8522 | 15.8163
2017-07-25 14:50 | 12908 | 8523 | 15.7996
2017-07-25 15:00 | 12910 | 8522 | 15.7963
2017-07-25 15:10 | 12907 | 8523 | 15.8013
2017-07-25 15:20 | 12904 | 8523 | 15.8063
2017-07-25 15:30 | 12898 | 8522 | 15.8163
2017-07-25 15:40 | 12889 | 8522 | 15.8314
2017-07-25 15:50 | 12874 | 8522 | 15.8566
2017-07-25 16:00 | 12856 | 8522 | 15.8868
2017-07-25 16:10 | 12838 | 8522 | 15.9171
2017-07-25 16:20 | 12813 | 8522 | 15.9592
2017-07-25 16:30 | 12797 | 8523 | 15.9863
2017-07-25 16:40 | 12784 | 8523 | 16.0082
2017-07-25 16:50 | 12771 | 8523 | 16.0302
2017-07-25 17:00 | 12761 | 8523 | 16.0472
2017-07-25 17:10 | 12749 | 8523 | 16.0675
2017-07-25 17:20 | 12738 | 8522 | 16.0862
2017-07-25 17:30 | 12728 | 8523 | 16.1032
2017-07-25 17:40 | 12718 | 8522 | 16.1202
2017-07-25 17:50 | 12710 | 8522 | 16.1338
2017-07-25 18:00 | 12703 | 8522 | 16.1457
2017-07-25 18:10 | 12697 | 8522 | 16.156
2017-07-25 18:20 | 12692 | 8522 | 16.1645
2017-07-25 18:30 | 12678 | 8522 | 16.1884
2017-07-25 18:40 | 12646 | 8522 | 16.2431
2017-07-25 18:50 | 12633 | 8522 | 16.2654
2017-07-25 19:00 | 12625 | 8522 | 16.2791
2017-07-25 19:10 | 12614 | 8523 | 16.298
2017-07-25 19:20 | 12605 | 8522 | 16.3134
2017-07-25 19:30 | 12596 | 8522 | 16.3289
2017-07-25 19:40 | 12586 | 8522 | 16.3461
2017-07-25 19:50 | 12580 | 8522 | 16.3564
2017-07-25 20:00 | 12579 | 8523 | 16.3582
2017-07-25 20:10 | 12568 | 8522 | 16.3771
2017-07-25 20:20 | 12541 | 8522 | 16.4237
2017-07-25 20:30 | 12496 | 8522 | 16.5017
2017-07-25 20:40 | 12448 | 8522 | 16.5852
2017-07-25 20:50 | 12416 | 8522 | 16.641
2017-07-25 21:00 | 12390 | 8522 | 16.6865
2017-07-25 21:10 | 12366 | 8523 | 16.7286
2017-07-25 21:20 | 12342 | 8523 | 16.7708
2017-07-25 21:30 | 12317 | 8522 | 16.8149
2017-07-25 21:40 | 12259 | 8522 | 16.9174
2017-07-25 21:50 | 12204 | 8522 | 17.0152
2017-07-25 22:00 | 12189 | 8522 | 17.042
2017-07-25 22:10 | 12161 | 8523 | 17.092
2017-07-25 22:20 | 12105 | 8522 | 17.1925
2017-07-25 22:30 | 12029 | 8522 | 17.3297
2017-07-25 22:40 | 11952 | 8522 | 17.4697
2017-07-25 22:50 | 11896 | 8522 | 17.5721
2017-07-25 23:00 | 11875 | 8522 | 17.6107
2017-07-25 23:10 | 11879 | 8522 | 17.6033
2017-07-25 23:20 | 11882 | 8522 | 17.5978
2017-07-25 23:30 | 11887 | 8522 | 17.5886
2017-07-25 23:40 | 11894 | 8522 | 17.5758
2017-07-25 23:50 | 11901 | 8522 | 17.5629
2017-07-26 | 11907 | 8522 | 17.5519
2017-07-26 00:10 | 11914 | 8522 | 17.5391
2017-07-26 00:20 | 11920 | 8523 | 17.5281
2017-07-26 00:30 | 11925 | 8522 | 17.519
2017-07-26 00:40 | 11930 | 8522 | 17.5098
2017-07-26 00:50 | 11935 | 8522 | 17.5007
2017-07-26 01:00 | 11942 | 8522 | 17.4879
2017-07-26 01:10 | 11950 | 8523 | 17.4733
2017-07-26 01:20 | 11961 | 8522 | 17.4532
2017-07-26 01:30 | 11970 | 8522 | 17.4368
2017-07-26 01:40 | 11979 | 8522 | 17.4205
2017-07-26 01:50 | 11986 | 8522 | 17.4077
2017-07-26 02:00 | 11991 | 8523 | 17.3986
2017-07-26 02:10 | 11996 | 8522 | 17.3895
2017-07-26 02:20 | 12000 | 8523 | 17.3823
2017-07-26 02:30 | 12003 | 8522 | 17.3768
2017-07-26 02:40 | 12007 | 8522 | 17.3696
2017-07-26 02:50 | 12009 | 8522 | 17.3659
2017-07-26 03:00 | 12008 | 8522 | 17.3678
2017-07-26 03:10 | 12005 | 8522 | 17.3732
2017-07-26 03:20 | 12001 | 8522 | 17.3805
2017-07-26 03:30 | 11996 | 8522 | 17.3895
2017-07-26 03:40 | 11986 | 8522 | 17.4077
2017-07-26 03:50 | 11974 | 8522 | 17.4296
2017-07-26 04:00 | 11965 | 8522 | 17.446
2017-07-26 04:10 | 11962 | 8522 | 17.4514
2017-07-26 04:20 | 11962 | 8522 | 17.4514
2017-07-26 04:30 | 11967 | 8522 | 17.4423
2017-07-26 04:40 | 11977 | 8522 | 17.4241
2017-07-26 04:50 | 11990 | 8522 | 17.4004
2017-07-26 05:00 | 11995 | 8522 | 17.3914
2017-07-26 05:10 | 12011 | 8522 | 17.3623
2017-07-26 05:20 | 12024 | 8522 | 17.3387
2017-07-26 05:30 | 12035 | 8522 | 17.3188
2017-07-26 05:40 | 12046 | 8522 | 17.2989
2017-07-26 05:50 | 12052 | 8522 | 17.2881
2017-07-26 06:00 | 12042 | 8522 | 17.3061
2017-07-26 06:10 | 12039 | 8522 | 17.3116
2017-07-26 06:20 | 12043 | 8522 | 17.3043
2017-07-26 06:30 | 12050 | 8522 | 17.2917
2017-07-26 06:40 | 12060 | 8522 | 17.2736
2017-07-26 06:50 | 12074 | 8522 | 17.2483
2017-07-26 07:00 | 12091 | 8522 | 17.2177
2017-07-26 07:10 | 12110 | 8522 | 17.1835
2017-07-26 07:20 | 12130 | 8522 | 17.1476
2017-07-26 07:30 | 12154 | 8522 | 17.1045
2017-07-26 07:40 | 12179 | 8522 | 17.0598
2017-07-26 07:50 | 12207 | 8522 | 17.0099
2017-07-26 08:00 | 12238 | 8523 | 16.9547
2017-07-26 08:10 | 12270 | 8522 | 16.8979
2017-07-26 08:20 | 12303 | 8522 | 16.8396
2017-07-26 08:30 | 12336 | 8522 | 16.7814
2017-07-26 08:40 | 12370 | 8522 | 16.7216
2017-07-26 08:50 | 12404 | 8522 | 16.662
2017-07-26 09:00 | 12436 | 8522 | 16.6061
2017-07-26 09:10 | 12468 | 8522 | 16.5503
2017-07-26 09:20 | 12500 | 8522 | 16.4947
2017-07-26 09:30 | 12532 | 8523 | 16.4393
2017-07-26 09:40 | 12563 | 8523 | 16.3857
2017-07-26 09:50 | 12594 | 8522 | 16.3323
2017-07-26 10:00 | 12624 | 8522 | 16.2808
2017-07-26 10:10 | 12655 | 8522 | 16.2277
2017-07-26 10:20 | 12685 | 8522 | 16.1764
2017-07-26 10:30 | 12716 | 8522 | 16.1236
2017-07-26 10:40 | 12747 | 8522 | 16.0709
2017-07-26 10:50 | 12778 | 8522 | 16.0184
2017-07-26 11:00 | 12808 | 8522 | 15.9677
2017-07-26 11:10 | 12837 | 8522 | 15.9188
2017-07-26 11:20 | 12865 | 8522 | 15.8717
2017-07-26 11:30 | 12892 | 8522 | 15.8264
2017-07-26 11:40 | 12916 | 8522 | 15.7862
2017-07-26 11:50 | 12938 | 8522 | 15.7495
2017-07-26 12:00 | 12958 | 8522 | 15.7161
2017-07-26 12:10 | 12978 | 8522 | 15.6828
2017-07-26 12:20 | 12999 | 8522 | 15.6479
2017-07-26 12:30 | 13019 | 8522 | 15.6148
2017-07-26 12:40 | 13039 | 8522 | 15.5816
2017-07-26 12:50 | 13059 | 8522 | 15.5486
2017-07-26 13:00 | 13078 | 8522 | 15.5172
2017-07-26 13:10 | 13098 | 8522 | 15.4843
2017-07-26 13:20 | 13118 | 8522 | 15.4514
2017-07-26 13:30 | 13138 | 8522 | 15.4186
2017-07-26 13:40 | 13150 | 8522 | 15.3989
2017-07-26 13:50 | 13159 | 8522 | 15.3841
2017-07-26 14:00 | 13172 | 8522 | 15.3629
2017-07-26 14:10 | 13188 | 8522 | 15.3367
2017-07-26 14:20 | 13201 | 8522 | 15.3155
2017-07-26 14:30 | 13213 | 8522 | 15.2959
2017-07-26 14:40 | 13222 | 8522 | 15.2813
2017-07-26 14:50 | 13225 | 8522 | 15.2764
2017-07-26 15:00 | 13219 | 8522 | 15.2861
2017-07-26 15:10 | 13212 | 8522 | 15.2976
2017-07-26 15:20 | 13207 | 8522 | 15.3057
2017-07-26 15:30 | 13203 | 8522 | 15.3122
2017-07-26 15:40 | 13194 | 8522 | 15.3269
2017-07-26 15:50 | 13180 | 8522 | 15.3498
2017-07-26 16:00 | 13162 | 8522 | 15.3792
2017-07-26 16:10 | 13143 | 8522 | 15.4104
2017-07-26 16:20 | 13122 | 8522 | 15.4448
2017-07-26 16:30 | 13100 | 8523 | 15.481
2017-07-26 16:40 | 13074 | 8522 | 15.5238
2017-07-26 16:50 | 13046 | 8522 | 15.5701
2017-07-26 17:00 | 13016 | 8522 | 15.6197
2017-07-26 17:10 | 12984 | 8522 | 15.6729
2017-07-26 17:20 | 12953 | 8522 | 15.7245
2017-07-26 17:30 | 12924 | 8522 | 15.7729
2017-07-26 17:40 | 12897 | 8522 | 15.818
2017-07-26 17:50 | 12872 | 8522 | 15.8599
2017-07-26 18:00 | 12843 | 8522 | 15.9087
2017-07-26 18:10 | 12811 | 8522 | 15.9626
2017-07-26 18:20 | 12774 | 8522 | 16.0252
2017-07-26 18:30 | 12741 | 8523 | 16.0811
2017-07-26 18:40 | 12711 | 8522 | 16.1321
2017-07-26 18:50 | 12683 | 8522 | 16.1798
2017-07-26 19:00 | 12658 | 8522 | 16.2226
2017-07-26 19:10 | 12633 | 8523 | 16.2654
2017-07-26 19:20 | 12607 | 8522 | 16.31
2017-07-26 19:30 | 12575 | 8522 | 16.3651
2017-07-26 19:40 | 12542 | 8522 | 16.422
2017-07-26 19:50 | 12492 | 8522 | 16.5086
2017-07-26 20:00 | 12429 | 8522 | 16.6183
2017-07-26 20:10 | 12378 | 8522 | 16.7076
2017-07-26 20:20 | 12337 | 8522 | 16.7796
2017-07-26 20:30 | 12298 | 8522 | 16.8484
2017-07-26 20:40 | 12262 | 8522 | 16.9121
2017-07-26 20:50 | 12214 | 8522 | 16.9974
2017-07-26 21:00 | 12130 | 8522 | 17.1476
2017-07-26 21:10 | 12039 | 8522 | 17.3116
2017-07-26 21:20 | 11874 | 8522 | 17.6125
2017-07-26 21:30 | 11716 | 8522 | 17.9051
2017-07-26 21:40 | 11568 | 8522 | 18.1833
2017-07-26 21:50 | 11429 | 8522 | 18.4482
2017-07-26 22:00 | 11304 | 8522 | 18.6896
2017-07-26 22:10 | 11212 | 8522 | 18.8692
2017-07-26 22:20 | 11169 | 8522 | 18.9537
2017-07-26 22:30 | 11168 | 8522 | 18.9556
2017-07-26 22:40 | 11182 | 8522 | 18.9281
2017-07-26 22:50 | 11201 | 8522 | 18.8907
2017-07-26 23:00 | 11212 | 8522 | 18.8692
2017-07-26 23:10 | 11200 | 8522 | 18.8927
2017-07-26 23:20 | 11189 | 8522 | 18.9143
2017-07-26 23:30 | 11180 | 8522 | 18.932
2017-07-26 23:40 | 11172 | 8522 | 18.9478
2017-07-26 23:50 | 11167 | 8522 | 18.9576
2017-07-27 | 11163 | 8522 | 18.9655
2017-07-27 00:10 | 11148 | 8522 | 18.9951
2017-07-27 00:20 | 11132 | 8522 | 19.0267
2017-07-27 00:30 | 11118 | 8522 | 19.0544
2017-07-27 00:40 | 11102 | 8522 | 19.0861
2017-07-27 00:50 | 11079 | 8522 | 19.1317
2017-07-27 01:00 | 11016 | 8522 | 19.2573
2017-07-27 01:10 | 10951 | 8522 | 19.3878
2017-07-27 01:20 | 10887 | 8522 | 19.5171
2017-07-27 01:30 | 10844 | 8522 | 19.6044
2017-07-27 01:40 | 10808 | 8522 | 19.6778
2017-07-27 01:50 | 10775 | 8522 | 19.7454
2017-07-27 02:00 | 10740 | 8522 | 19.8173
2017-07-27 02:10 | 10655 | 8522 | 19.993
2017-07-27 02:20 | 10528 | 8522 | 20.2586
2017-07-27 02:30 | 10398 | 8522 | 20.5341
2017-07-27 02:40 | 10270 | 8522 | 20.8093
2017-07-27 02:50 | 10179 | 8522 | 21.0073
2017-07-27 03:00 | 10083 | 8522 | 21.2183
2017-07-27 03:10 | 10038 | 8522 | 21.318
2017-07-27 03:20 | 10055 | 8522 | 21.2803
2017-07-27 03:30 | 10093 | 8522 | 21.1962
2017-07-27 03:40 | 10135 | 8522 | 21.1037
2017-07-27 03:50 | 10174 | 8522 | 21.0182
2017-07-27 04:00 | 10209 | 8522 | 20.9418
2017-07-27 04:10 | 10239 | 8522 | 20.8765
2017-07-27 04:20 | 10265 | 8522 | 20.8201
2017-07-27 04:30 | 10286 | 8522 | 20.7747
2017-07-27 04:40 | 10305 | 8522 | 20.7337
2017-07-27 04:50 | 10321 | 8522 | 20.6992
2017-07-27 05:00 | 10332 | 8522 | 20.6755
2017-07-27 05:10 | 10328 | 8522 | 20.6841
2017-07-27 05:20 | 10303 | 8522 | 20.738
2017-07-27 05:30 | 10259 | 8522 | 20.8331
2017-07-27 05:40 | 10183 | 8522 | 20.9985
2017-07-27 05:50 | 10143 | 8522 | 21.0861
2017-07-27 06:00 | 10133 | 8522 | 21.1081
2017-07-27 06:10 | 10145 | 8522 | 21.0817
2017-07-27 06:20 | 10171 | 8522 | 21.0248
2017-07-27 06:30 | 10206 | 8522 | 20.9483
2017-07-27 06:40 | 10249 | 8522 | 20.8548
2017-07-27 06:50 | 10296 | 8522 | 20.7531
2017-07-27 07:00 | 10347 | 8522 | 20.6433
2017-07-27 07:10 | 10400 | 8522 | 20.5299
2017-07-27 07:20 | 10454 | 8522 | 20.415
2017-07-27 07:30 | 10510 | 8522 | 20.2965
2017-07-27 07:40 | 10567 | 8522 | 20.1766
2017-07-27 07:50 | 10623 | 8522 | 20.0596
2017-07-27 08:00 | 10679 | 8522 | 19.9433
2017-07-27 08:10 | 10734 | 8522 | 19.8297
2017-07-27 08:20 | 10786 | 8522 | 19.7229
2017-07-27 08:30 | 10837 | 8522 | 19.6187
2017-07-27 08:40 | 10886 | 8522 | 19.5191
2017-07-27 08:50 | 10933 | 8522 | 19.424
2017-07-27 09:00 | 10978 | 8522 | 19.3335
2017-07-27 09:10 | 11021 | 8522 | 19.2473
2017-07-27 09:20 | 11063 | 8522 | 19.1635
2017-07-27 09:30 | 11104 | 8522 | 19.0821
2017-07-27 09:40 | 11144 | 8522 | 19.003
2017-07-27 09:50 | 11184 | 8522 | 18.9241
2017-07-27 10:00 | 11222 | 8522 | 18.8496
2017-07-27 10:10 | 11259 | 8522 | 18.7772
2017-07-27 10:20 | 11294 | 8522 | 18.709
2017-07-27 10:30 | 11328 | 8522 | 18.643
2017-07-27 10:40 | 11361 | 8522 | 18.5792
2017-07-27 10:50 | 11394 | 8522 | 18.5155
2017-07-27 11:00 | 11427 | 8522 | 18.4521
2017-07-27 11:10 | 11458 | 8522 | 18.3927
2017-07-27 11:20 | 11489 | 8522 | 18.3334
2017-07-27 11:30 | 11520 | 8522 | 18.2744
2017-07-27 11:40 | 11550 | 8522 | 18.2174
2017-07-27 11:50 | 11579 | 8522 | 18.1625
2017-07-27 12:00 | 11608 | 8522 | 18.1077
2017-07-27 12:10 | 11636 | 8522 | 18.055
2017-07-27 12:20 | 11661 | 8522 | 18.008
2017-07-27 12:30 | 11683 | 8523 | 17.9668
2017-07-27 12:40 | 11704 | 8522 | 17.9275
2017-07-27 12:50 | 11724 | 8522 | 17.8902
2017-07-27 13:00 | 11743 | 8522 | 17.8548
2017-07-27 13:10 | 11761 | 8522 | 17.8213
2017-07-27 13:20 | 11778 | 8522 | 17.7898
2017-07-27 13:30 | 11795 | 8522 | 17.7583
2017-07-27 13:40 | 11811 | 8522 | 17.7287
2017-07-27 13:50 | 11826 | 8522 | 17.7009
2017-07-27 14:00 | 11842 | 8522 | 17.6714
2017-07-27 14:10 | 11858 | 8522 | 17.6419
2017-07-27 14:20 | 11874 | 8523 | 17.6125
2017-07-27 14:30 | 11891 | 8523 | 17.5813
2017-07-27 14:40 | 11907 | 8522 | 17.5519
2017-07-27 14:50 | 11919 | 8522 | 17.53
2017-07-27 15:00 | 11926 | 8522 | 17.5172
2017-07-27 15:10 | 11930 | 8522 | 17.5098
2017-07-27 15:20 | 11933 | 8522 | 17.5044
2017-07-27 15:30 | 11933 | 8522 | 17.5044
2017-07-27 15:40 | 11930 | 8523 | 17.5098
2017-07-27 15:50 | 11922 | 8522 | 17.5245
2017-07-27 16:00 | 11916 | 8522 | 17.5355
2017-07-27 16:10 | 11914 | 8522 | 17.5391
2017-07-27 16:20 | 11910 | 8522 | 17.5465
2017-07-27 16:30 | 11904 | 8522 | 17.5574
2017-07-27 16:40 | 11890 | 8522 | 17.5831
2017-07-27 16:50 | 11870 | 8522 | 17.6199
2017-07-27 17:00 | 11855 | 8522 | 17.6475
2017-07-27 17:10 | 11848 | 8522 | 17.6604
2017-07-27 17:20 | 11845 | 8522 | 17.6659
2017-07-27 17:30 | 11844 | 8523 | 17.6677
2017-07-27 17:40 | 11840 | 8523 | 17.6751
2017-07-27 17:50 | 11835 | 8522 | 17.6843
2017-07-27 18:00 | 11829 | 8522 | 17.6954
2017-07-27 18:10 | 11822 | 8523 | 17.7083
2017-07-27 18:20 | 11814 | 8522 | 17.7231
2017-07-27 18:30 | 11805 | 8522 | 17.7398
2017-07-27 18:40 | 11795 | 8522 | 17.7583
2017-07-27 18:50 | 11785 | 8522 | 17.7768
2017-07-27 19:00 | 11776 | 8522 | 17.7935
2017-07-27 19:10 | 11767 | 8522 | 17.8102
2017-07-27 19:20 | 11755 | 8522 | 17.8325
2017-07-27 19:30 | 11732 | 8522 | 17.8753
2017-07-27 19:40 | 11700 | 8522 | 17.935
2017-07-27 19:50 | 11662 | 8522 | 18.0062
2017-07-27 20:00 | 11616 | 8522 | 18.0927
2017-07-27 20:10 | 11513 | 8522 | 18.2877
2017-07-27 20:20 | 11387 | 8522 | 18.529
2017-07-27 20:30 | 11249 | 8522 | 18.7967
2017-07-27 20:40 | 11112 | 8522 | 19.0662
2017-07-27 20:50 | 10993 | 8522 | 19.3034
2017-07-27 21:00 | 10947 | 8522 | 19.3958
2017-07-27 21:10 | 10947 | 8522 | 19.3958
2017-07-27 21:20 | 10969 | 8522 | 19.3516
2017-07-27 21:30 | 11004 | 8522 | 19.2813
2017-07-27 21:40 | 11044 | 8522 | 19.2014
2017-07-27 21:50 | 11085 | 8522 | 19.1198
2017-07-27 22:00 | 11127 | 8522 | 19.0365
2017-07-27 22:10 | 11166 | 8522 | 18.9596
2017-07-27 22:20 | 11200 | 8522 | 18.8927
2017-07-27 22:30 | 11230 | 8522 | 18.8339
2017-07-27 22:40 | 11257 | 8522 | 18.7811
2017-07-27 22:50 | 11282 | 8522 | 18.7324
2017-07-27 23:00 | 11303 | 8522 | 18.6915
2017-07-27 23:10 | 11326 | 8522 | 18.6469
2017-07-27 23:20 | 11343 | 8522 | 18.614
2017-07-27 23:30 | 11349 | 8522 | 18.6024
2017-07-27 23:40 | 11331 | 8522 | 18.6372
2017-07-27 23:50 | 11296 | 8522 | 18.7051
2017-07-28 | 11185 | 8522 | 18.9222
2017-07-28 00:10 | 11017 | 8522 | 19.2553
2017-07-28 00:20 | 10862 | 8522 | 19.5678
2017-07-28 00:30 | 10753 | 8522 | 19.7906
2017-07-28 00:40 | 10650 | 8522 | 20.0034
2017-07-28 00:50 | 10552 | 8522 | 20.2081
2017-07-28 01:00 | 10488 | 8522 | 20.3429
2017-07-28 01:10 | 10417 | 8522 | 20.4936
2017-07-28 01:20 | 10393 | 8522 | 20.5448
2017-07-28 01:30 | 10436 | 8522 | 20.4532
2017-07-28 01:40 | 10502 | 8522 | 20.3134
2017-07-28 01:50 | 10570 | 8522 | 20.1704
2017-07-28 02:00 | 10631 | 8522 | 20.0429
2017-07-28 02:10 | 10665 | 8522 | 19.9723
2017-07-28 02:20 | 10697 | 8522 | 19.906
2017-07-28 02:30 | 10724 | 8522 | 19.8503
2017-07-28 02:40 | 10744 | 8522 | 19.8091
2017-07-28 02:50 | 10752 | 8522 | 19.7926
2017-07-28 03:00 | 10756 | 8522 | 19.7844
2017-07-28 03:10 | 10763 | 8522 | 19.77
2017-07-28 03:20 | 10771 | 8522 | 19.7536
2017-07-28 03:30 | 10733 | 8522 | 19.8317
2017-07-28 03:40 | 10697 | 8522 | 19.906
2017-07-28 03:50 | 10676 | 8522 | 19.9495
2017-07-28 04:00 | 10658 | 8522 | 19.9868
2017-07-28 04:10 | 10640 | 8522 | 20.0242
2017-07-28 04:20 | 10621 | 8522 | 20.0638
2017-07-28 04:30 | 10602 | 8522 | 20.1034
2017-07-28 04:40 | 10583 | 8522 | 20.1431
2017-07-28 04:50 | 10564 | 8522 | 20.1829
2017-07-28 05:00 | 10544 | 8522 | 20.2249
2017-07-28 05:10 | 10537 | 8522 | 20.2396
2017-07-28 05:20 | 10545 | 8522 | 20.2228
2017-07-28 05:30 | 10562 | 8522 | 20.1871
2017-07-28 05:40 | 10582 | 8522 | 20.1452
2017-07-28 05:50 | 10606 | 8522 | 20.0951
2017-07-28 06:00 | 10632 | 8522 | 20.0409
2017-07-28 06:10 | 10656 | 8522 | 19.991
2017-07-28 06:20 | 10680 | 8522 | 19.9412
2017-07-28 06:30 | 10702 | 8522 | 19.8957
2017-07-28 06:40 | 10724 | 8522 | 19.8503
2017-07-28 06:50 | 10747 | 8522 | 19.8029
2017-07-28 07:00 | 10771 | 8522 | 19.7536
2017-07-28 07:10 | 10795 | 8522 | 19.7044
2017-07-28 07:20 | 10819 | 8522 | 19.6554
2017-07-28 07:30 | 10843 | 8522 | 19.6064
2017-07-28 07:40 | 10865 | 8522 | 19.5617
2017-07-28 07:50 | 10888 | 8522 | 19.515
2017-07-28 08:00 | 10910 | 8522 | 19.4705
2017-07-28 08:10 | 10933 | 8522 | 19.424
2017-07-28 08:20 | 10957 | 8522 | 19.3757
2017-07-28 08:30 | 10980 | 8522 | 19.3295
2017-07-28 08:40 | 11002 | 8522 | 19.2853
2017-07-28 08:50 | 11023 | 8522 | 19.2433
2017-07-28 09:00 | 11043 | 8522 | 19.2034
2017-07-28 09:10 | 11063 | 8522 | 19.1635
2017-07-28 09:20 | 11082 | 8522 | 19.1257
2017-07-28 09:30 | 11101 | 8522 | 19.088
2017-07-28 09:40 | 11119 | 8522 | 19.0524
2017-07-28 09:50 | 11137 | 8522 | 19.0168
2017-07-28 10:00 | 11157 | 8522 | 18.9773
2017-07-28 10:10 | 11176 | 8522 | 18.9399
2017-07-28 10:20 | 11196 | 8522 | 18.9006
2017-07-28 10:30 | 11217 | 8522 | 18.8594
2017-07-28 10:40 | 11239 | 8522 | 18.8163
2017-07-28 10:50 | 11261 | 8522 | 18.7733
2017-07-28 11:00 | 11282 | 8522 | 18.7324
2017-07-28 11:10 | 11305 | 8522 | 18.6876
2017-07-28 11:20 | 11328 | 8522 | 18.643
2017-07-28 11:30 | 11351 | 8522 | 18.5985
2017-07-28 11:40 | 11374 | 8522 | 18.5541
2017-07-28 11:50 | 11398 | 8522 | 18.5078
2017-07-28 12:00 | 11421 | 8522 | 18.4636
2017-07-28 12:10 | 11442 | 8522 | 18.4233
2017-07-28 12:20 | 11462 | 8522 | 18.385
2017-07-28 12:30 | 11481 | 8522 | 18.3487
2017-07-28 12:40 | 11501 | 8522 | 18.3106
2017-07-28 12:50 | 11520 | 8522 | 18.2744
2017-07-28 13:00 | 11538 | 8522 | 18.2402
2017-07-28 13:10 | 11555 | 8522 | 18.2079
2017-07-28 13:20 | 11572 | 8522 | 18.1757
2017-07-28 13:30 | 11589 | 8522 | 18.1436
2017-07-28 13:40 | 11606 | 8523 | 18.1115
2017-07-28 13:50 | 11623 | 8522 | 18.0795
2017-07-28 14:00 | 11638 | 8522 | 18.0512
2017-07-28 14:10 | 11654 | 8522 | 18.0212
2017-07-28 14:20 | 11669 | 8522 | 17.993
2017-07-28 14:30 | 11679 | 8522 | 17.9743
2017-07-28 14:40 | 11688 | 8522 | 17.9575
2017-07-28 14:50 | 11697 | 8522 | 17.9406
2017-07-28 15:00 | 11704 | 8522 | 17.9275
2017-07-28 15:10 | 11709 | 8522 | 17.9182
2017-07-28 15:20 | 11713 | 8522 | 17.9107
2017-07-28 15:30 | 11715 | 8522 | 17.907
2017-07-28 15:40 | 11713 | 8522 | 17.9107
2017-07-28 15:50 | 11709 | 8522 | 17.9182
2017-07-28 16:00 | 11704 | 8522 | 17.9275
2017-07-28 16:10 | 11697 | 8522 | 17.9406
2017-07-28 16:20 | 11691 | 8522 | 17.9518
2017-07-28 16:30 | 11683 | 8522 | 17.9668
2017-07-28 16:40 | 11647 | 8522 | 18.0343
2017-07-28 16:50 | 11606 | 8522 | 18.1115
2017-07-28 17:00 | 11566 | 8522 | 18.1871
2017-07-28 17:10 | 11543 | 8522 | 18.2307
2017-07-28 17:20 | 11531 | 8522 | 18.2535
2017-07-28 17:30 | 11524 | 8522 | 18.2668
2017-07-28 17:40 | 11519 | 8522 | 18.2763
2017-07-28 17:50 | 11515 | 8522 | 18.2839
2017-07-28 18:00 | 11511 | 8522 | 18.2915
2017-07-28 18:10 | 11506 | 8522 | 18.301
2017-07-28 18:20 | 11501 | 8522 | 18.3106
2017-07-28 18:30 | 11495 | 8522 | 18.322
2017-07-28 18:40 | 11489 | 8522 | 18.3334
2017-07-28 18:50 | 11481 | 8522 | 18.3487
2017-07-28 19:00 | 11472 | 8522 | 18.3659
2017-07-28 19:10 | 11461 | 8522 | 18.3869
2017-07-28 19:20 | 11452 | 8522 | 18.4041
2017-07-28 19:30 | 11443 | 8522 | 18.4214
2017-07-28 19:40 | 11433 | 8522 | 18.4406
2017-07-28 19:50 | 11424 | 8522 | 18.4578
2017-07-28 20:00 | 11410 | 8522 | 18.4847
2017-07-28 20:10 | 11385 | 8522 | 18.5328
2017-07-28 20:20 | 11212 | 8522 | 18.8692
2017-07-28 20:30 | 11720 | 8522 | 17.8977
2017-07-28 20:40 | 12085 | 8522 | 17.2285
2017-07-28 20:50 | 12397 | 8522 | 16.6743
2017-07-28 21:00 | 12652 | 8522 | 16.2328
2017-07-28 21:10 | 12846 | 8522 | 15.9036
2017-07-28 21:20 | 12982 | 8522 | 15.6762
2017-07-28 21:30 | 13080 | 8522 | 15.5139
2017-07-28 21:40 | 13168 | 8522 | 15.3694
2017-07-28 21:50 | 13248 | 8523 | 15.239
2017-07-28 22:00 | 13317 | 8522 | 15.1272
2017-07-28 22:10 | 13378 | 8522 | 15.0289
2017-07-28 22:20 | 13431 | 8522 | 14.9439
2017-07-28 22:30 | 13477 | 8523 | 14.8704
2017-07-28 22:40 | 13529 | 8522 | 14.7877
2017-07-28 22:50 | 13576 | 8522 | 14.7132
2017-07-28 23:00 | 13294 | 8522 | 15.1644
2017-07-28 23:10 | 13389 | 8522 | 15.0112
2017-07-28 23:20 | 13113 | 8522 | 15.4596
2017-07-28 23:30 | 13389 | 8522 | 15.0112
2017-07-28 23:40 | 13800 | 8523 | 14.3623
2017-07-28 23:50 | 13897 | 8523 | 14.2124
2017-07-29 | 13903 | 8522 | 14.2031
2017-07-29 00:10 | 13995 | 8523 | 14.0621
2017-07-29 00:20 | 14001 | 8522 | 14.0529
2017-07-29 00:30 | 14090 | 8523 | 13.9175
2017-07-29 00:40 | 14266 | 8522 | 13.6524
2017-07-29 00:50 | 14552 | 8522 | 13.2296
2017-07-29 01:00 | 14761 | 8522 | 12.9264
2017-07-29 01:10 | 14650 | 8522 | 13.0868
2017-07-29 01:20 | 14571 | 8522 | 13.2018
2017-07-29 01:30 | 14555 | 8522 | 13.2252
2017-07-29 01:40 | 14644 | 8522 | 13.0955
2017-07-29 01:50 | 14710 | 8522 | 12.9999
2017-07-29 02:00 | 14694 | 8522 | 13.023
2017-07-29 02:10 | 14587 | 8522 | 13.1784
2017-07-29 02:20 | 14354 | 8522 | 13.5213
2017-07-29 02:30 | 14221 | 8522 | 13.7199
2017-07-29 02:40 | 14145 | 8522 | 13.8342
2017-07-29 02:50 | 14142 | 8522 | 13.8388
2017-07-29 03:00 | 14076 | 8522 | 13.9387
2017-07-29 03:10 | 13961 | 8522 | 14.1141
2017-07-29 03:20 | 13852 | 8522 | 14.2818
2017-07-29 03:30 | 13703 | 8522 | 14.5135
2017-07-29 03:40 | 13621 | 8522 | 14.6422
2017-07-29 03:50 | 13575 | 8522 | 14.7148
2017-07-29 04:00 | 13849 | 8522 | 14.2864
2017-07-29 04:10 | 14062 | 8522 | 13.96
2017-07-29 04:20 | 13729 | 8523 | 14.4729
2017-07-29 04:30 | 13279 | 8522 | 15.1887
2017-07-29 04:40 | 13365 | 8522 | 15.0498
2017-07-29 04:50 | 13657 | 8522 | 14.5856
2017-07-29 05:00 | 14066 | 8523 | 13.9539
2017-07-29 05:10 | 14133 | 8522 | 13.8524
2017-07-29 05:20 | 14226 | 8522 | 13.7124
2017-07-29 05:30 | 14249 | 8522 | 13.6779
2017-07-29 05:40 | 14832 | 8522 | 12.8245
2017-07-29 05:50 | 14984 | 8522 | 12.6082
2017-07-29 06:00 | 15421 | 8522 | 11.9998
2017-07-29 06:10 | 15688 | 8522 | 11.6375
2017-07-29 06:20 | 15707 | 8522 | 11.612
2017-07-29 06:30 | 19170 | 8522 | 7.4631
2017-07-29 06:40 | 21036 | 8523 | 5.5629

## Sheet2
2017-07-14 20:10 | 8313 | 8313 | 8313 | 8313 | 8313 | 8521 | 0
2017-07-14 20:20 | 8450 | 8450 | 8450 | 8450 | 8450 | 8521 | 0
2017-07-14 20:30 | 8482 | 8482 | 8482 | 8482 | 8482 | 8521 | 0
2017-07-14 20:40 | 8095 | 8095 | 8095 | 8095 | 8095 | 8521 | 0
2017-07-14 20:50 | 8147 | 8147 | 8147 | 8147 | 8147 | 8521 | 0
2017-07-14 21:00 | 8143 | 8143 | 8143 | 8143 | 8143 | 8521 | 0
2017-07-14 21:10 | 8221 | 8221 | 8221 | 8221 | 8221 | 8521 | 0
2017-07-14 21:20 | 8187 | 8187 | 8187 | 8187 | 8187 | 8521 | 0
2017-07-14 21:30 | 8064 | 8064 | 8064 | 8064 | 8064 | 8521 | 0
2017-07-14 21:40 | 7954 | 7954 | 7954 | 7954 | 7954 | 8521 | 0
2017-07-14 21:50 | 7865 | 7865 | 7865 | 7865 | 7865 | 8521 | 0
2017-07-14 22:00 | 7765 | 7765 | 7765 | 7765 | 7765 | 8521 | 0
2017-07-14 22:10 | 7660 | 7660 | 7660 | 7660 | 7660 | 8521 | 0
2017-07-14 22:20 | 7564 | 7564 | 7564 | 7564 | 7564 | 8520 | 0
2017-07-14 22:30 | 7479 | 7479 | 7479 | 7479 | 7479 | 8520 | 0
2017-07-14 22:40 | 7406 | 7406 | 7406 | 7406 | 7406 | 8520 | 0
2017-07-14 22:50 | 7345 | 7345 | 7345 | 7345 | 7345 | 8520 | 0
2017-07-14 23:00 | 7269 | 7269 | 7269 | 7269 | 7269 | 8520 | 0
2017-07-14 23:10 | 7189 | 7189 | 7189 | 7189 | 7189 | 8520 | 0
2017-07-14 23:20 | 7126 | 7126 | 7126 | 7126 | 7126 | 8520 | 0
2017-07-14 23:30 | 7083 | 7083 | 7083 | 7083 | 7083 | 8520 | 0
2017-07-14 23:40 | 7069 | 7069 | 7069 | 7069 | 7069 | 8520 | 0
2017-07-14 23:50 | 7195 | 7195 | 7195 | 7195 | 7195 | 8520 | 0
2017-07-15 | 7337 | 7337 | 7337 | 7337 | 7337 | 8520 | 0
2017-07-15 00:10 | 7482 | 7482 | 7482 | 7482 | 7482 | 8520 | 0
2017-07-15 00:20 | 7631 | 7631 | 7631 | 7631 | 7631 | 8520 | 0
2017-07-15 00:30 | 7780 | 7780 | 7780 | 7780 | 7780 | 8521 | 0
2017-07-15 00:40 | 7927 | 7927 | 7927 | 7927 | 7927 | 8521 | 0
2017-07-15 00:50 | 8069 | 8069 | 8069 | 8069 | 8069 | 8521 | 0
2017-07-15 01:00 | 8205 | 8205 | 8205 | 8205 | 8205 | 8521 | 0
2017-07-15 01:10 | 8333 | 8333 | 8333 | 8333 | 8333 | 8521 | 0
2017-07-15 01:20 | 8453 | 8453 | 8453 | 8453 | 8453 | 8521 | 0
2017-07-15 01:30 | 8565 | 8565 | 8565 | 8565 | 8565 | 8521 | 0
2017-07-15 01:40 | 8668 | 8668 | 8668 | 8668 | 8668 | 8521 | 0
2017-07-15 01:50 | 8765 | 8765 | 8765 | 8765 | 8765 | 8521 | 0
2017-07-15 02:00 | 8854 | 8854 | 8854 | 8854 | 8854 | 8521 | 0
2017-07-15 02:10 | 8938 | 8938 | 8938 | 8938 | 8938 | 8522 | 0
2017-07-15 02:20 | 9015 | 9015 | 9015 | 9015 | 9015 | 8522 | 0
2017-07-15 02:30 | 9087 | 9087 | 9087 | 9087 | 9087 | 8522 | 0
2017-07-15 02:40 | 9154 | 9154 | 9154 | 9154 | 9154 | 8522 | 0
2017-07-15 02:50 | 9216 | 9216 | 9216 | 9216 | 9216 | 8522 | 0
2017-07-15 03:00 | 9274 | 9274 | 9274 | 9274 | 9274 | 8522 | 0
2017-07-15 03:10 | 9328 | 9328 | 9328 | 9328 | 9328 | 8522 | 0
2017-07-15 03:20 | 9378 | 9378 | 9378 | 9378 | 9378 | 8522 | 0
2017-07-15 03:30 | 9424 | 9424 | 9424 | 9424 | 9424 | 8522 | 0
2017-07-15 03:40 | 9466 | 9466 | 9466 | 9466 | 9466 | 8522 | 0
2017-07-15 03:50 | 9506 | 9506 | 9506 | 9506 | 9506 | 8522 | 0
2017-07-15 04:00 | 9543 | 9543 | 9543 | 9543 | 9543 | 8522 | 0
2017-07-15 04:10 | 9578 | 9578 | 9578 | 9578 | 9578 | 8522 | 0
2017-07-15 04:20 | 9610 | 9610 | 9610 | 9610 | 9610 | 8522 | 0
2017-07-15 04:30 | 9641 | 9641 | 9641 | 9641 | 9641 | 8522 | 0
2017-07-15 04:40 | 9669 | 9669 | 9669 | 9669 | 9669 | 8522 | 0
2017-07-15 04:50 | 9695 | 9695 | 9695 | 9695 | 9695 | 8522 | 0
2017-07-15 05:00 | 9719 | 9719 | 9719 | 9719 | 9719 | 8522 | 0
2017-07-15 05:10 | 9742 | 9742 | 9742 | 9742 | 9742 | 8522 | 0
2017-07-15 05:20 | 9764 | 9764 | 9764 | 9764 | 9764 | 8522 | 0
2017-07-15 05:30 | 9786 | 9786 | 9786 | 9786 | 9786 | 8522 | 0
2017-07-15 05:40 | 9809 | 9809 | 9809 | 9809 | 9809 | 8522 | 0
2017-07-15 05:50 | 9833 | 9833 | 9833 | 9833 | 9833 | 8522 | 0
2017-07-15 06:00 | 9856 | 9856 | 9856 | 9856 | 9856 | 8522 | 0
2017-07-15 06:10 | 9880 | 9880 | 9880 | 9880 | 9880 | 8522 | 0
2017-07-15 06:20 | 9904 | 9904 | 9904 | 9904 | 9904 | 8522 | 0
2017-07-15 06:30 | 9927 | 9927 | 9927 | 9927 | 9927 | 8522 | 0
2017-07-15 06:40 | 9951 | 9951 | 9951 | 9951 | 9951 | 8522 | 0
2017-07-15 06:50 | 9974 | 9974 | 9974 | 9974 | 9974 | 8522 | 0
2017-07-15 07:00 | 9997 | 9997 | 9997 | 9997 | 9997 | 8522 | 0
2017-07-15 07:10 | 10020 | 10020 | 10020 | 10020 | 10020 | 8522 | 0
2017-07-15 07:20 | 10043 | 10043 | 10043 | 10043 | 10043 | 8522 | 0
2017-07-15 07:30 | 10066 | 10066 | 10066 | 10066 | 10066 | 8522 | 0
2017-07-15 07:40 | 10089 | 10089 | 10089 | 10089 | 10089 | 8522 | 0
2017-07-15 07:50 | 10112 | 10112 | 10112 | 10112 | 10112 | 8522 | 0
2017-07-15 08:00 | 10136 | 10136 | 10136 | 10136 | 10136 | 8522 | 0
2017-07-15 08:10 | 10159 | 10159 | 10159 | 10159 | 10159 | 8522 | 0
2017-07-15 08:20 | 10181 | 10181 | 10181 | 10181 | 10181 | 8522 | 0
2017-07-15 08:30 | 10205 | 10205 | 10205 | 10205 | 10205 | 8522 | 0
2017-07-15 08:40 | 10228 | 10228 | 10228 | 10228 | 10228 | 8522 | 0
2017-07-15 08:50 | 10251 | 10251 | 10251 | 10251 | 10251 | 8522 | 0
2017-07-15 09:00 | 10274 | 10274 | 10274 | 10274 | 10274 | 8522 | 0
2017-07-15 09:10 | 10297 | 10297 | 10297 | 10297 | 10297 | 8522 | 0
2017-07-15 09:20 | 10321 | 10321 | 10321 | 10321 | 10321 | 8522 | 0
2017-07-15 09:30 | 10344 | 10344 | 10344 | 10344 | 10344 | 8522 | 0
2017-07-15 09:40 | 10368 | 10368 | 10368 | 10368 | 10368 | 8522 | 0
2017-07-15 09:50 | 10391 | 10391 | 10391 | 10391 | 10391 | 8522 | 0
2017-07-15 10:00 | 10414 | 10414 | 10414 | 10414 | 10414 | 8522 | 0
2017-07-15 10:10 | 10438 | 10438 | 10438 | 10438 | 10438 | 8522 | 0
2017-07-15 10:20 | 10462 | 10462 | 10462 | 10462 | 10462 | 8522 | 0
2017-07-15 10:30 | 10486 | 10486 | 10486 | 10486 | 10486 | 8522 | 0
2017-07-15 10:40 | 10510 | 10510 | 10510 | 10510 | 10510 | 8522 | 0
2017-07-15 10:50 | 10535 | 10535 | 10535 | 10535 | 10535 | 8522 | 0
2017-07-15 11:00 | 10560 | 10560 | 10560 | 10560 | 10560 | 8522 | 0
2017-07-15 11:10 | 10586 | 10586 | 10586 | 10586 | 10586 | 8522 | 0
2017-07-15 11:20 | 10610 | 10610 | 10610 | 10610 | 10610 | 8522 | 0
2017-07-15 11:30 | 10635 | 10635 | 10635 | 10635 | 10635 | 8522 | 0
2017-07-15 11:40 | 10659 | 10659 | 10659 | 10659 | 10659 | 8522 | 0
2017-07-15 11:50 | 10683 | 10683 | 10683 | 10683 | 10683 | 8522 | 0
2017-07-15 12:00 | 10706 | 10706 | 10706 | 10706 | 10706 | 8522 | 0
2017-07-15 12:10 | 10728 | 10728 | 10728 | 10728 | 10728 | 8522 | 0
2017-07-15 12:20 | 10750 | 10750 | 10750 | 10750 | 10750 | 8522 | 0
2017-07-15 12:30 | 10771 | 10771 | 10771 | 10771 | 10771 | 8522 | 0
2017-07-15 12:40 | 10793 | 10793 | 10793 | 10793 | 10793 | 8522 | 0
2017-07-15 12:50 | 10814 | 10814 | 10814 | 10814 | 10814 | 8522 | 0
2017-07-15 13:00 | 10835 | 10835 | 10835 | 10835 | 10835 | 8522 | 0
2017-07-15 13:10 | 10856 | 10856 | 10856 | 10856 | 10856 | 8522 | 0
2017-07-15 13:20 | 10877 | 10877 | 10877 | 10877 | 10877 | 8522 | 0
2017-07-15 13:30 | 10898 | 10898 | 10898 | 10898 | 10898 | 8522 | 0
2017-07-15 13:40 | 10918 | 10918 | 10918 | 10918 | 10918 | 8522 | 0
2017-07-15 13:50 | 10937 | 10937 | 10937 | 10937 | 10937 | 8522 | 0
2017-07-15 14:00 | 10955 | 10955 | 10955 | 10955 | 10955 | 8522 | 0
2017-07-15 14:10 | 10973 | 10973 | 10973 | 10973 | 10973 | 8522 | 0
2017-07-15 14:20 | 10990 | 10990 | 10990 | 10990 | 10990 | 8522 | 0
2017-07-15 14:30 | 11006 | 11006 | 11006 | 11006 | 11006 | 8522 | 0
2017-07-15 14:40 | 11022 | 11022 | 11022 | 11022 | 11022 | 8522 | 0
2017-07-15 14:50 | 11038 | 11038 | 11038 | 11038 | 11038 | 8522 | 0
2017-07-15 15:00 | 11054 | 11054 | 11054 | 11054 | 11054 | 8522 | 0
2017-07-15 15:10 | 11069 | 11069 | 11069 | 11069 | 11069 | 8522 | 0
2017-07-15 15:20 | 11084 | 11084 | 11084 | 11084 | 11084 | 8522 | 0
2017-07-15 15:30 | 11097 | 11097 | 11097 | 11097 | 11097 | 8523 | 0
2017-07-15 15:40 | 11109 | 11109 | 11109 | 11109 | 11109 | 8522 | 0
2017-07-15 15:50 | 11121 | 11121 | 11121 | 11121 | 11121 | 8522 | 0
2017-07-15 16:00 | 11132 | 11132 | 11132 | 11132 | 11132 | 8522 | 0
2017-07-15 16:10 | 11141 | 11141 | 11141 | 11141 | 11141 | 8522 | 0
2017-07-15 16:20 | 11150 | 11150 | 11150 | 11150 | 11150 | 8522 | 0
2017-07-15 16:30 | 11158 | 11158 | 11158 | 11158 | 11158 | 8522 | 0
2017-07-15 16:40 | 11164 | 11164 | 11164 | 11164 | 11164 | 8522 | 0
2017-07-15 16:50 | 11170 | 11170 | 11170 | 11170 | 11170 | 8522 | 0
2017-07-15 17:00 | 11175 | 11175 | 11175 | 11175 | 11175 | 8523 | 0
2017-07-15 17:10 | 11179 | 11179 | 11179 | 11179 | 11179 | 8522 | 0
2017-07-15 17:20 | 11182 | 11182 | 11182 | 11182 | 11182 | 8522 | 0
2017-07-15 17:30 | 11184 | 11184 | 11184 | 11184 | 11184 | 8523 | 0
2017-07-15 17:40 | 11186 | 11186 | 11186 | 11186 | 11186 | 8522 | 0
2017-07-15 17:50 | 11187 | 11187 | 11187 | 11187 | 11187 | 8522 | 0
2017-07-15 18:00 | 11186 | 11186 | 11186 | 11186 | 11186 | 8522 | 0
2017-07-15 18:10 | 11184 | 11184 | 11184 | 11184 | 11184 | 8522 | 0
2017-07-15 18:20 | 11180 | 11180 | 11180 | 11180 | 11180 | 8522 | 0
2017-07-15 18:30 | 11176 | 11176 | 11176 | 11176 | 11176 | 8522 | 0
2017-07-15 18:40 | 11170 | 11170 | 11170 | 11170 | 11170 | 8522 | 0
2017-07-15 18:50 | 11164 | 11164 | 11164 | 11164 | 11164 | 8522 | 0
2017-07-15 19:00 | 11157 | 11157 | 11157 | 11157 | 11157 | 8522 | 0
2017-07-15 19:10 | 11149 | 11149 | 11149 | 11149 | 11149 | 8522 | 0
2017-07-15 19:20 | 11141 | 11141 | 11141 | 11141 | 11141 | 8522 | 0
2017-07-15 19:30 | 11131 | 11131 | 11131 | 11131 | 11131 | 8522 | 0
2017-07-15 19:40 | 11122 | 11122 | 11122 | 11122 | 11122 | 8522 | 0
2017-07-15 19:50 | 11112 | 11112 | 11112 | 11112 | 11112 | 8522 | 0
2017-07-15 20:00 | 11102 | 11102 | 11102 | 11102 | 11102 | 8522 | 0
2017-07-15 20:10 | 11092 | 11092 | 11092 | 11092 | 11092 | 8522 | 0
2017-07-15 20:20 | 11082 | 11082 | 11082 | 11082 | 11082 | 8522 | 0
2017-07-15 20:30 | 11072 | 11072 | 11072 | 11072 | 11072 | 8522 | 0
2017-07-15 20:40 | 11062 | 11062 | 11062 | 11062 | 11062 | 8522 | 0
2017-07-15 20:50 | 11052 | 11052 | 11052 | 11052 | 11052 | 8522 | 0
2017-07-15 21:00 | 11042 | 11042 | 11042 | 11042 | 11042 | 8522 | 0
2017-07-15 21:10 | 11031 | 11031 | 11031 | 11031 | 11031 | 8522 | 0
2017-07-15 21:20 | 11019 | 11019 | 11019 | 11019 | 11019 | 8522 | 0
2017-07-15 21:30 | 11005 | 11005 | 11005 | 11005 | 11005 | 8522 | 0
2017-07-15 21:40 | 10991 | 10991 | 10991 | 10991 | 10991 | 8522 | 0
2017-07-15 21:50 | 10977 | 10977 | 10977 | 10977 | 10977 | 8522 | 0
2017-07-15 22:00 | 10961 | 10961 | 10961 | 10961 | 10961 | 8522 | 0
2017-07-15 22:10 | 10945 | 10945 | 10945 | 10945 | 10945 | 8522 | 0
2017-07-15 22:20 | 10928 | 10928 | 10928 | 10928 | 10928 | 8522 | 0
2017-07-15 22:30 | 10912 | 10912 | 10912 | 10912 | 10912 | 8522 | 0
2017-07-15 22:40 | 10894 | 10894 | 10894 | 10894 | 10894 | 8522 | 0
2017-07-15 22:50 | 10876 | 10876 | 10876 | 10876 | 10876 | 8522 | 0
2017-07-15 23:00 | 10856 | 10856 | 10856 | 10856 | 10856 | 8522 | 0
2017-07-15 23:10 | 10836 | 10836 | 10836 | 10836 | 10836 | 8522 | 0
2017-07-15 23:20 | 10718 | 10718 | 10718 | 10718 | 10718 | 8522 | 0
2017-07-15 23:30 | 10485 | 10485 | 10485 | 10485 | 10485 | 8522 | 0
2017-07-15 23:40 | 10384 | 10384 | 10384 | 10384 | 10384 | 8522 | 0
2017-07-15 23:50 | 10296 | 10296 | 10296 | 10296 | 10296 | 8522 | 0
2017-07-16 | 10239 | 10239 | 10239 | 10239 | 10239 | 8522 | 0
2017-07-16 00:10 | 10207 | 10207 | 10207 | 10207 | 10207 | 8522 | 0
2017-07-16 00:20 | 10183 | 10183 | 10183 | 10183 | 10183 | 8522 | 0
2017-07-16 00:30 | 10164 | 10164 | 10164 | 10164 | 10164 | 8522 | 0
2017-07-16 00:40 | 10151 | 10151 | 10151 | 10151 | 10151 | 8522 | 0
2017-07-16 00:50 | 10141 | 10141 | 10141 | 10141 | 10141 | 8522 | 0
2017-07-16 01:00 | 10135 | 10135 | 10135 | 10135 | 10135 | 8522 | 0
2017-07-16 01:10 | 10132 | 10132 | 10132 | 10132 | 10132 | 8522 | 0
2017-07-16 01:20 | 10131 | 10131 | 10131 | 10131 | 10131 | 8522 | 0
2017-07-16 01:30 | 10132 | 10132 | 10132 | 10132 | 10132 | 8522 | 0
2017-07-16 01:40 | 10134 | 10134 | 10134 | 10134 | 10134 | 8522 | 0
2017-07-16 01:50 | 10134 | 10134 | 10134 | 10134 | 10134 | 8522 | 0
2017-07-16 02:00 | 10121 | 10121 | 10121 | 10121 | 10121 | 8522 | 0
2017-07-16 02:10 | 10105 | 10105 | 10105 | 10105 | 10105 | 8522 | 0
2017-07-16 02:20 | 10087 | 10087 | 10087 | 10087 | 10087 | 8522 | 0
2017-07-16 02:30 | 10069 | 10069 | 10069 | 10069 | 10069 | 8522 | 0
2017-07-16 02:40 | 10033 | 10033 | 10033 | 10033 | 10033 | 8522 | 0
2017-07-16 02:50 | 10001 | 10001 | 10001 | 10001 | 10001 | 8522 | 0
2017-07-16 03:00 | 9993 | 9993 | 9993 | 9993 | 9993 | 8522 | 0
2017-07-16 03:10 | 9988 | 9988 | 9988 | 9988 | 9988 | 8522 | 0
2017-07-16 03:20 | 9982 | 9982 | 9982 | 9982 | 9982 | 8522 | 0
2017-07-16 03:30 | 9977 | 9977 | 9977 | 9977 | 9977 | 8522 | 0
2017-07-16 03:40 | 9962 | 9962 | 9962 | 9962 | 9962 | 8522 | 0
2017-07-16 03:50 | 9764 | 9764 | 9764 | 9764 | 9764 | 8522 | 0
2017-07-16 04:00 | 9635 | 9635 | 9635 | 9635 | 9635 | 8522 | 0
2017-07-16 04:10 | 9574 | 9574 | 9574 | 9574 | 9574 | 8522 | 0
2017-07-16 04:20 | 9544 | 9544 | 9544 | 9544 | 9544 | 8522 | 0
2017-07-16 04:30 | 9528 | 9528 | 9528 | 9528 | 9528 | 8522 | 0
2017-07-16 04:40 | 9512 | 9512 | 9512 | 9512 | 9512 | 8522 | 0
2017-07-16 04:50 | 9504 | 9504 | 9504 | 9504 | 9504 | 8522 | 0
2017-07-16 05:00 | 9507 | 9507 | 9507 | 9507 | 9507 | 8522 | 0
2017-07-16 05:10 | 9518 | 9518 | 9518 | 9518 | 9518 | 8522 | 0
2017-07-16 05:20 | 9533 | 9533 | 9533 | 9533 | 9533 | 8522 | 0
2017-07-16 05:30 | 9551 | 9551 | 9551 | 9551 | 9551 | 8522 | 0
2017-07-16 05:40 | 9572 | 9572 | 9572 | 9572 | 9572 | 8522 | 0
2017-07-16 05:50 | 9595 | 9595 | 9595 | 9595 | 9595 | 8522 | 0
2017-07-16 06:00 | 9618 | 9618 | 9618 | 9618 | 9618 | 8522 | 0
2017-07-16 06:10 | 9642 | 9642 | 9642 | 9642 | 9642 | 8522 | 0
2017-07-16 06:20 | 9667 | 9667 | 9667 | 9667 | 9667 | 8522 | 0
2017-07-16 06:30 | 9691 | 9691 | 9691 | 9691 | 9691 | 8522 | 0
2017-07-16 06:40 | 9717 | 9717 | 9717 | 9717 | 9717 | 8522 | 0
2017-07-16 06:50 | 9742 | 9742 | 9742 | 9742 | 9742 | 8522 | 0
2017-07-16 07:00 | 9767 | 9767 | 9767 | 9767 | 9767 | 8522 | 0
2017-07-16 07:10 | 9793 | 9793 | 9793 | 9793 | 9793 | 8522 | 0
2017-07-16 07:20 | 9818 | 9818 | 9818 | 9818 | 9818 | 8522 | 0
2017-07-16 07:30 | 9843 | 9843 | 9843 | 9843 | 9843 | 8522 | 0
2017-07-16 07:40 | 9867 | 9867 | 9867 | 9867 | 9867 | 8522 | 0
2017-07-16 07:50 | 9892 | 9892 | 9892 | 9892 | 9892 | 8522 | 0
2017-07-16 08:00 | 9916 | 9916 | 9916 | 9916 | 9916 | 8522 | 0
2017-07-16 08:10 | 9940 | 9940 | 9940 | 9940 | 9940 | 8522 | 0
2017-07-16 08:20 | 9963 | 9963 | 9963 | 9963 | 9963 | 8522 | 0
2017-07-16 08:30 | 9986 | 9986 | 9986 | 9986 | 9986 | 8522 | 0
2017-07-16 08:40 | 10009 | 10009 | 10009 | 10009 | 10009 | 8522 | 0
2017-07-16 08:50 | 10032 | 10032 | 10032 | 10032 | 10032 | 8522 | 0
2017-07-16 09:00 | 10054 | 10054 | 10054 | 10054 | 10054 | 8522 | 0
2017-07-16 09:10 | 10077 | 10077 | 10077 | 10077 | 10077 | 8522 | 0
2017-07-16 09:20 | 10099 | 10099 | 10099 | 10099 | 10099 | 8522 | 0
2017-07-16 09:30 | 10122 | 10122 | 10122 | 10122 | 10122 | 8522 | 0
2017-07-16 09:40 | 10145 | 10145 | 10145 | 10145 | 10145 | 8522 | 0
2017-07-16 09:50 | 10167 | 10167 | 10167 | 10167 | 10167 | 8522 | 0
2017-07-16 10:00 | 10189 | 10189 | 10189 | 10189 | 10189 | 8522 | 0
2017-07-16 10:10 | 10211 | 10211 | 10211 | 10211 | 10211 | 8522 | 0
2017-07-16 10:20 | 10232 | 10232 | 10232 | 10232 | 10232 | 8522 | 0
2017-07-16 10:30 | 10254 | 10254 | 10254 | 10254 | 10254 | 8522 | 0
2017-07-16 10:40 | 10275 | 10275 | 10275 | 10275 | 10275 | 8522 | 0
2017-07-16 10:50 | 10296 | 10296 | 10296 | 10296 | 10296 | 8522 | 0
2017-07-16 11:00 | 10317 | 10317 | 10317 | 10317 | 10317 | 8522 | 0
2017-07-16 11:10 | 10337 | 10337 | 10337 | 10337 | 10337 | 8522 | 0
2017-07-16 11:20 | 10357 | 10357 | 10357 | 10357 | 10357 | 8522 | 0
2017-07-16 11:30 | 10377 | 10377 | 10377 | 10377 | 10377 | 8522 | 0
2017-07-16 11:40 | 10396 | 10396 | 10396 | 10396 | 10396 | 8522 | 0
2017-07-16 11:50 | 10415 | 10415 | 10415 | 10415 | 10415 | 8522 | 0
2017-07-16 12:00 | 10434 | 10434 | 10434 | 10434 | 10434 | 8522 | 0
2017-07-16 12:10 | 10453 | 10453 | 10453 | 10453 | 10453 | 8522 | 0
2017-07-16 12:20 | 10472 | 10472 | 10472 | 10472 | 10472 | 8522 | 0
2017-07-16 12:30 | 10491 | 10491 | 10491 | 10491 | 10491 | 8522 | 0
2017-07-16 12:40 | 10509 | 10509 | 10509 | 10509 | 10509 | 8522 | 0
2017-07-16 12:50 | 10527 | 10527 | 10527 | 10527 | 10527 | 8522 | 0
2017-07-16 13:00 | 10545 | 10545 | 10545 | 10545 | 10545 | 8522 | 0
2017-07-16 13:10 | 10563 | 10563 | 10563 | 10563 | 10563 | 8522 | 0
2017-07-16 13:20 | 10581 | 10581 | 10581 | 10581 | 10581 | 8522 | 0
2017-07-16 13:30 | 10598 | 10598 | 10598 | 10598 | 10598 | 8522 | 0
2017-07-16 13:40 | 10616 | 10616 | 10616 | 10616 | 10616 | 8522 | 0
2017-07-16 13:50 | 10633 | 10633 | 10633 | 10633 | 10633 | 8522 | 0
2017-07-16 14:00 | 10649 | 10649 | 10649 | 10649 | 10649 | 8522 | 0
2017-07-16 14:10 | 10665 | 10665 | 10665 | 10665 | 10665 | 8522 | 0
2017-07-16 14:20 | 10667 | 10667 | 10667 | 10667 | 10667 | 8522 | 0
2017-07-16 14:30 | 10656 | 10656 | 10656 | 10656 | 10656 | 8522 | 0
2017-07-16 14:40 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-16 14:50 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-16 15:00 | 10654 | 10654 | 10654 | 10654 | 10654 | 8522 | 0
2017-07-16 15:10 | 10658 | 10658 | 10658 | 10658 | 10658 | 8522 | 0
2017-07-16 15:20 | 10671 | 10671 | 10671 | 10671 | 10671 | 8522 | 0
2017-07-16 15:30 | 10682 | 10682 | 10682 | 10682 | 10682 | 8522 | 0
2017-07-16 15:40 | 10691 | 10691 | 10691 | 10691 | 10691 | 8522 | 0
2017-07-16 15:50 | 10700 | 10700 | 10700 | 10700 | 10700 | 8522 | 0
2017-07-16 16:00 | 10711 | 10711 | 10711 | 10711 | 10711 | 8522 | 0
2017-07-16 16:10 | 10722 | 10722 | 10722 | 10722 | 10722 | 8522 | 0
2017-07-16 16:20 | 10730 | 10730 | 10730 | 10730 | 10730 | 8522 | 0
2017-07-16 16:30 | 10734 | 10734 | 10734 | 10734 | 10734 | 8522 | 0
2017-07-16 16:40 | 10735 | 10735 | 10735 | 10735 | 10735 | 8522 | 0
2017-07-16 16:50 | 10734 | 10734 | 10734 | 10734 | 10734 | 8522 | 0
2017-07-16 17:00 | 10706 | 10706 | 10706 | 10706 | 10706 | 8522 | 0
2017-07-16 17:10 | 10698 | 10698 | 10698 | 10698 | 10698 | 8522 | 0
2017-07-16 17:20 | 10685 | 10685 | 10685 | 10685 | 10685 | 8522 | 0
2017-07-16 17:30 | 10666 | 10666 | 10666 | 10666 | 10666 | 8522 | 0
2017-07-16 17:40 | 10653 | 10653 | 10653 | 10653 | 10653 | 8522 | 0
2017-07-16 17:50 | 10640 | 10640 | 10640 | 10640 | 10640 | 8522 | 0
2017-07-16 18:00 | 10630 | 10630 | 10630 | 10630 | 10630 | 8522 | 0
2017-07-16 18:10 | 10618 | 10618 | 10618 | 10618 | 10618 | 8522 | 0
2017-07-16 18:20 | 10600 | 10600 | 10600 | 10600 | 10600 | 8522 | 0
2017-07-16 18:30 | 10568 | 10568 | 10568 | 10568 | 10568 | 8522 | 0
2017-07-16 18:40 | 10519 | 10519 | 10519 | 10519 | 10519 | 8522 | 0
2017-07-16 18:50 | 10461 | 10461 | 10461 | 10461 | 10461 | 8522 | 0
2017-07-16 19:00 | 10395 | 10395 | 10395 | 10395 | 10395 | 8522 | 0
2017-07-16 19:10 | 10322 | 10322 | 10322 | 10322 | 10322 | 8522 | 0
2017-07-16 19:20 | 10244 | 10244 | 10244 | 10244 | 10244 | 8522 | 0
2017-07-16 19:30 | 10157 | 10157 | 10157 | 10157 | 10157 | 8522 | 0
2017-07-16 19:40 | 10065 | 10065 | 10065 | 10065 | 10065 | 8522 | 0
2017-07-16 19:50 | 9966 | 9966 | 9966 | 9966 | 9966 | 8522 | 0
2017-07-16 20:00 | 9862 | 9862 | 9862 | 9862 | 9862 | 8522 | 0
2017-07-16 20:10 | 9756 | 9756 | 9756 | 9756 | 9756 | 8522 | 0
2017-07-16 20:20 | 9653 | 9653 | 9653 | 9653 | 9653 | 8522 | 0
2017-07-16 20:30 | 9552 | 9552 | 9552 | 9552 | 9552 | 8522 | 0
2017-07-16 20:40 | 9444 | 9444 | 9444 | 9444 | 9444 | 8522 | 0
2017-07-16 20:50 | 9337 | 9337 | 9337 | 9337 | 9337 | 8522 | 0
2017-07-16 21:00 | 9245 | 9245 | 9245 | 9245 | 9245 | 8522 | 0
2017-07-16 21:10 | 9108 | 9108 | 9108 | 9108 | 9108 | 8522 | 0
2017-07-16 21:20 | 9000 | 9000 | 9000 | 9000 | 9000 | 8522 | 0
2017-07-16 21:30 | 8905 | 8905 | 8905 | 8905 | 8905 | 8522 | 0
2017-07-16 21:40 | 8817 | 8817 | 8817 | 8817 | 8817 | 8522 | 0
2017-07-16 21:50 | 8735 | 8735 | 8735 | 8735 | 8735 | 8522 | 0
2017-07-16 22:00 | 8655 | 8655 | 8655 | 8655 | 8655 | 8521 | 0
2017-07-16 22:10 | 8577 | 8577 | 8577 | 8577 | 8577 | 8521 | 0
2017-07-16 22:20 | 8501 | 8501 | 8501 | 8501 | 8501 | 8521 | 0
2017-07-16 22:30 | 8427 | 8427 | 8427 | 8427 | 8427 | 8521 | 0
2017-07-16 22:40 | 8355 | 8355 | 8355 | 8355 | 8355 | 8521 | 0
2017-07-16 22:50 | 8286 | 8286 | 8286 | 8286 | 8286 | 8521 | 0
2017-07-16 23:00 | 8219 | 8219 | 8219 | 8219 | 8219 | 8521 | 0
2017-07-16 23:10 | 8155 | 8155 | 8155 | 8155 | 8155 | 8521 | 0
2017-07-16 23:20 | 8093 | 8093 | 8093 | 8093 | 8093 | 8521 | 0
2017-07-16 23:30 | 8033 | 8033 | 8033 | 8033 | 8033 | 8521 | 0
2017-07-16 23:40 | 7977 | 7977 | 7977 | 7977 | 7977 | 8521 | 0
2017-07-16 23:50 | 7922 | 7922 | 7922 | 7922 | 7922 | 8521 | 0
2017-07-17 | 7870 | 7870 | 7870 | 7870 | 7870 | 8521 | 0
2017-07-17 00:10 | 7821 | 7821 | 7821 | 7821 | 7821 | 8521 | 0
2017-07-17 00:20 | 7776 | 7776 | 7776 | 7776 | 7776 | 8521 | 0
2017-07-17 00:30 | 7732 | 7732 | 7732 | 7732 | 7732 | 8521 | 0
2017-07-17 00:40 | 7692 | 7692 | 7692 | 7692 | 7692 | 8521 | 0
2017-07-17 00:50 | 7654 | 7654 | 7654 | 7654 | 7654 | 8520 | 0
2017-07-17 01:00 | 7620 | 7620 | 7620 | 7620 | 7620 | 8520 | 0
2017-07-17 01:10 | 7587 | 7587 | 7587 | 7587 | 7587 | 8520 | 0
2017-07-17 01:20 | 7557 | 7557 | 7557 | 7557 | 7557 | 8520 | 0
2017-07-17 01:30 | 7530 | 7530 | 7530 | 7530 | 7530 | 8520 | 0
2017-07-17 01:40 | 7505 | 7505 | 7505 | 7505 | 7505 | 8520 | 0
2017-07-17 01:50 | 7482 | 7482 | 7482 | 7482 | 7482 | 8520 | 0
2017-07-17 02:00 | 7462 | 7462 | 7462 | 7462 | 7462 | 8520 | 0
2017-07-17 02:10 | 7444 | 7444 | 7444 | 7444 | 7444 | 8520 | 0
2017-07-17 02:20 | 7428 | 7428 | 7428 | 7428 | 7428 | 8520 | 0
2017-07-17 02:30 | 7415 | 7415 | 7415 | 7415 | 7415 | 8520 | 0
2017-07-17 02:40 | 7403 | 7403 | 7403 | 7403 | 7403 | 8520 | 0
2017-07-17 02:50 | 7394 | 7394 | 7394 | 7394 | 7394 | 8520 | 0
2017-07-17 03:00 | 7387 | 7387 | 7387 | 7387 | 7387 | 8520 | 0
2017-07-17 03:10 | 7382 | 7382 | 7382 | 7382 | 7382 | 8520 | 0
2017-07-17 03:20 | 7379 | 7379 | 7379 | 7379 | 7379 | 8520 | 0
2017-07-17 03:30 | 7379 | 7379 | 7379 | 7379 | 7379 | 8520 | 0
2017-07-17 03:40 | 7381 | 7381 | 7381 | 7381 | 7381 | 8520 | 0
2017-07-17 03:50 | 7385 | 7385 | 7385 | 7385 | 7385 | 8520 | 0
2017-07-17 04:00 | 7391 | 7391 | 7391 | 7391 | 7391 | 8520 | 0
2017-07-17 04:10 | 7400 | 7400 | 7400 | 7400 | 7400 | 8520 | 0
2017-07-17 04:20 | 7410 | 7410 | 7410 | 7410 | 7410 | 8520 | 0
2017-07-17 04:30 | 7422 | 7422 | 7422 | 7422 | 7422 | 8520 | 0
2017-07-17 04:40 | 7436 | 7436 | 7436 | 7436 | 7436 | 8520 | 0
2017-07-17 04:50 | 7452 | 7452 | 7452 | 7452 | 7452 | 8520 | 0
2017-07-17 05:00 | 7470 | 7470 | 7470 | 7470 | 7470 | 8520 | 0
2017-07-17 05:10 | 7491 | 7491 | 7491 | 7491 | 7491 | 8520 | 0
2017-07-17 05:20 | 7515 | 7515 | 7515 | 7515 | 7515 | 8520 | 0
2017-07-17 05:30 | 7541 | 7541 | 7541 | 7541 | 7541 | 8520 | 0
2017-07-17 05:40 | 7570 | 7570 | 7570 | 7570 | 7570 | 8520 | 0
2017-07-17 05:50 | 7603 | 7603 | 7603 | 7603 | 7603 | 8520 | 0
2017-07-17 06:00 | 7639 | 7639 | 7639 | 7639 | 7639 | 8520 | 0
2017-07-17 06:10 | 7679 | 7679 | 7679 | 7679 | 7679 | 8520 | 0
2017-07-17 06:20 | 7722 | 7722 | 7722 | 7722 | 7722 | 8520 | 0
2017-07-17 06:30 | 7767 | 7767 | 7767 | 7767 | 7767 | 8521 | 0
2017-07-17 06:40 | 7816 | 7816 | 7816 | 7816 | 7816 | 8521 | 0
2017-07-17 06:50 | 7868 | 7868 | 7868 | 7868 | 7868 | 8521 | 0
2017-07-17 07:00 | 7922 | 7922 | 7922 | 7922 | 7922 | 8521 | 0
2017-07-17 07:10 | 7979 | 7979 | 7979 | 7979 | 7979 | 8521 | 0
2017-07-17 07:20 | 8038 | 8038 | 8038 | 8038 | 8038 | 8521 | 0
2017-07-17 07:30 | 8100 | 8100 | 8100 | 8100 | 8100 | 8521 | 0
2017-07-17 07:40 | 8165 | 8165 | 8165 | 8165 | 8165 | 8521 | 0
2017-07-17 07:50 | 8231 | 8231 | 8231 | 8231 | 8231 | 8521 | 0
2017-07-17 08:00 | 8299 | 8299 | 8299 | 8299 | 8299 | 8521 | 0
2017-07-17 08:10 | 8369 | 8369 | 8369 | 8369 | 8369 | 8521 | 0
2017-07-17 08:20 | 8441 | 8441 | 8441 | 8441 | 8441 | 8521 | 0
2017-07-17 08:30 | 8514 | 8514 | 8514 | 8514 | 8514 | 8521 | 0
2017-07-17 08:40 | 8588 | 8588 | 8588 | 8588 | 8588 | 8521 | 0
2017-07-17 08:50 | 8664 | 8664 | 8664 | 8664 | 8664 | 8521 | 0
2017-07-17 09:00 | 8741 | 8741 | 8741 | 8741 | 8741 | 8521 | 0
2017-07-17 09:10 | 8819 | 8819 | 8819 | 8819 | 8819 | 8521 | 0
2017-07-17 09:20 | 8897 | 8897 | 8897 | 8897 | 8897 | 8522 | 0
2017-07-17 09:30 | 8977 | 8977 | 8977 | 8977 | 8977 | 8522 | 0
2017-07-17 09:40 | 9057 | 9057 | 9057 | 9057 | 9057 | 8522 | 0
2017-07-17 09:50 | 9137 | 9137 | 9137 | 9137 | 9137 | 8522 | 0
2017-07-17 10:00 | 9218 | 9218 | 9218 | 9218 | 9218 | 8522 | 0
2017-07-17 10:10 | 9299 | 9299 | 9299 | 9299 | 9299 | 8522 | 0
2017-07-17 10:20 | 9381 | 9381 | 9381 | 9381 | 9381 | 8522 | 0
2017-07-17 10:30 | 9464 | 9464 | 9464 | 9464 | 9464 | 8522 | 0
2017-07-17 10:40 | 9546 | 9546 | 9546 | 9546 | 9546 | 8522 | 0
2017-07-17 10:50 | 9629 | 9629 | 9629 | 9629 | 9629 | 8522 | 0
2017-07-17 11:00 | 9713 | 9713 | 9713 | 9713 | 9713 | 8522 | 0
2017-07-17 11:10 | 9797 | 9797 | 9797 | 9797 | 9797 | 8522 | 0
2017-07-17 11:20 | 9881 | 9881 | 9881 | 9881 | 9881 | 8522 | 0
2017-07-17 11:30 | 9966 | 9966 | 9966 | 9966 | 9966 | 8522 | 0
2017-07-17 11:40 | 10050 | 10050 | 10050 | 10050 | 10050 | 8522 | 0
2017-07-17 11:50 | 10135 | 10135 | 10135 | 10135 | 10135 | 8522 | 0
2017-07-17 12:00 | 10219 | 10219 | 10219 | 10219 | 10219 | 8522 | 0
2017-07-17 12:10 | 10303 | 10303 | 10303 | 10303 | 10303 | 8522 | 0
2017-07-17 12:20 | 10373 | 10373 | 10373 | 10373 | 10373 | 8522 | 0
2017-07-17 12:30 | 10437 | 10437 | 10437 | 10437 | 10437 | 8522 | 0
2017-07-17 12:40 | 10505 | 10505 | 10505 | 10505 | 10505 | 8522 | 0
2017-07-17 12:50 | 10574 | 10574 | 10574 | 10574 | 10574 | 8522 | 0
2017-07-17 13:00 | 10643 | 10643 | 10643 | 10643 | 10643 | 8522 | 0
2017-07-17 13:10 | 10713 | 10713 | 10713 | 10713 | 10713 | 8522 | 0
2017-07-17 13:20 | 10781 | 10781 | 10781 | 10781 | 10781 | 8522 | 0
2017-07-17 13:30 | 10846 | 10846 | 10846 | 10846 | 10846 | 8522 | 0
2017-07-17 13:40 | 10913 | 10913 | 10913 | 10913 | 10913 | 8522 | 0
2017-07-17 13:50 | 10978 | 10978 | 10978 | 10978 | 10978 | 8522 | 0
2017-07-17 14:00 | 11029 | 11029 | 11029 | 11029 | 11029 | 8522 | 0
2017-07-17 14:10 | 11074 | 11074 | 11074 | 11074 | 11074 | 8522 | 0
2017-07-17 14:20 | 11113 | 11113 | 11113 | 11113 | 11113 | 8522 | 0
2017-07-17 14:30 | 11144 | 11144 | 11144 | 11144 | 11144 | 8522 | 0
2017-07-17 14:40 | 11165 | 11165 | 11165 | 11165 | 11165 | 8522 | 0
2017-07-17 14:50 | 11172 | 11172 | 11172 | 11172 | 11172 | 8522 | 0
2017-07-17 15:00 | 11166 | 11166 | 11166 | 11166 | 11166 | 8522 | 0
2017-07-17 15:10 | 11155 | 11155 | 11155 | 11155 | 11155 | 8522 | 0
2017-07-17 15:20 | 11146 | 11146 | 11146 | 11146 | 11146 | 8522 | 0
2017-07-17 15:30 | 11137 | 11137 | 11137 | 11137 | 11137 | 8522 | 0
2017-07-17 15:40 | 11119 | 11119 | 11119 | 11119 | 11119 | 8522 | 0
2017-07-17 15:50 | 11093 | 11093 | 11093 | 11093 | 11093 | 8522 | 0
2017-07-17 16:00 | 11069 | 11069 | 11069 | 11069 | 11069 | 8522 | 0
2017-07-17 16:10 | 11038 | 11038 | 11038 | 11038 | 11038 | 8522 | 0
2017-07-17 16:20 | 11002 | 11002 | 11002 | 11002 | 11002 | 8522 | 0
2017-07-17 16:30 | 10961 | 10961 | 10961 | 10961 | 10961 | 8522 | 0
2017-07-17 16:40 | 10911 | 10911 | 10911 | 10911 | 10911 | 8522 | 0
2017-07-17 16:50 | 10855 | 10855 | 10855 | 10855 | 10855 | 8522 | 0
2017-07-17 17:00 | 10786 | 10786 | 10786 | 10786 | 10786 | 8522 | 0
2017-07-17 17:10 | 10719 | 10719 | 10719 | 10719 | 10719 | 8522 | 0
2017-07-17 17:20 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-17 17:30 | 10568 | 10568 | 10568 | 10568 | 10568 | 8522 | 0
2017-07-17 17:40 | 10501 | 10501 | 10501 | 10501 | 10501 | 8522 | 0
2017-07-17 17:50 | 10435 | 10435 | 10435 | 10435 | 10435 | 8522 | 0
2017-07-17 18:00 | 10358 | 10358 | 10358 | 10358 | 10358 | 8522 | 0
2017-07-17 18:10 | 10276 | 10276 | 10276 | 10276 | 10276 | 8522 | 0
2017-07-17 18:20 | 10180 | 10180 | 10180 | 10180 | 10180 | 8522 | 0
2017-07-17 18:30 | 10077 | 10077 | 10077 | 10077 | 10077 | 8522 | 0
2017-07-17 18:40 | 9976 | 9976 | 9976 | 9976 | 9976 | 8522 | 0
2017-07-17 18:50 | 9881 | 9881 | 9881 | 9881 | 9881 | 8522 | 0
2017-07-17 19:00 | 9792 | 9792 | 9792 | 9792 | 9792 | 8522 | 0
2017-07-17 19:10 | 9705 | 9705 | 9705 | 9705 | 9705 | 8522 | 0
2017-07-17 19:20 | 9624 | 9624 | 9624 | 9624 | 9624 | 8522 | 0
2017-07-17 19:30 | 9542 | 9542 | 9542 | 9542 | 9542 | 8522 | 0
2017-07-17 19:40 | 9460 | 9460 | 9460 | 9460 | 9460 | 8522 | 0
2017-07-17 19:50 | 9378 | 9378 | 9378 | 9378 | 9378 | 8522 | 0
2017-07-17 20:00 | 9292 | 9292 | 9292 | 9292 | 9292 | 8522 | 0
2017-07-17 20:10 | 9202 | 9202 | 9202 | 9202 | 9202 | 8522 | 0
2017-07-17 20:20 | 9107 | 9107 | 9107 | 9107 | 9107 | 8522 | 0
2017-07-17 20:30 | 9011 | 9011 | 9011 | 9011 | 9011 | 8522 | 0
2017-07-17 20:40 | 8915 | 8915 | 8915 | 8915 | 8915 | 8522 | 0
2017-07-17 20:50 | 8816 | 8816 | 8816 | 8816 | 8816 | 8522 | 0
2017-07-17 21:00 | 8715 | 8715 | 8715 | 8715 | 8715 | 8522 | 0
2017-07-17 21:10 | 8614 | 8614 | 8614 | 8614 | 8614 | 8521 | 0
2017-07-17 21:20 | 8514 | 8514 | 8514 | 8514 | 8514 | 8521 | 0
2017-07-17 21:30 | 8410 | 8410 | 8410 | 8410 | 8410 | 8521 | 0
2017-07-17 21:40 | 8308 | 8308 | 8308 | 8308 | 8308 | 8521 | 0
2017-07-17 21:50 | 8210 | 8210 | 8210 | 8210 | 8210 | 8521 | 0
2017-07-17 22:00 | 8114 | 8114 | 8114 | 8114 | 8114 | 8521 | 0
2017-07-17 22:10 | 8023 | 8023 | 8023 | 8023 | 8023 | 8521 | 0
2017-07-17 22:20 | 7938 | 7938 | 7938 | 7938 | 7938 | 8521 | 0
2017-07-17 22:30 | 7860 | 7860 | 7860 | 7860 | 7860 | 8521 | 0
2017-07-17 22:40 | 7789 | 7789 | 7789 | 7789 | 7789 | 8521 | 0
2017-07-17 22:50 | 7723 | 7723 | 7723 | 7723 | 7723 | 8521 | 0
2017-07-17 23:00 | 7661 | 7661 | 7661 | 7661 | 7661 | 8521 | 0
2017-07-17 23:10 | 7605 | 7605 | 7605 | 7605 | 7605 | 8520 | 0
2017-07-17 23:20 | 7554 | 7554 | 7554 | 7554 | 7554 | 8520 | 0
2017-07-17 23:30 | 7507 | 7507 | 7507 | 7507 | 7507 | 8520 | 0
2017-07-17 23:40 | 7467 | 7467 | 7467 | 7467 | 7467 | 8520 | 0
2017-07-17 23:50 | 7437 | 7437 | 7437 | 7437 | 7437 | 8520 | 0
2017-07-18 | 7412 | 7412 | 7412 | 7412 | 7412 | 8520 | 0
2017-07-18 00:10 | 7387 | 7387 | 7387 | 7387 | 7387 | 8520 | 0
2017-07-18 00:20 | 7360 | 7360 | 7360 | 7360 | 7360 | 8520 | 0
2017-07-18 00:30 | 7338 | 7338 | 7338 | 7338 | 7338 | 8520 | 0
2017-07-18 00:40 | 7317 | 7317 | 7317 | 7317 | 7317 | 8520 | 0
2017-07-18 00:50 | 7293 | 7293 | 7293 | 7293 | 7293 | 8520 | 0
2017-07-18 01:00 | 7266 | 7266 | 7266 | 7266 | 7266 | 8520 | 0
2017-07-18 01:10 | 7237 | 7237 | 7237 | 7237 | 7237 | 8520 | 0
2017-07-18 01:20 | 7207 | 7207 | 7207 | 7207 | 7207 | 8520 | 0
2017-07-18 01:30 | 7179 | 7179 | 7179 | 7179 | 7179 | 8520 | 0
2017-07-18 01:40 | 7152 | 7152 | 7152 | 7152 | 7152 | 8520 | 0
2017-07-18 01:50 | 7125 | 7125 | 7125 | 7125 | 7125 | 8520 | 0
2017-07-18 02:00 | 7101 | 7101 | 7101 | 7101 | 7101 | 8520 | 0
2017-07-18 02:10 | 7085 | 7085 | 7085 | 7085 | 7085 | 8520 | 0
2017-07-18 02:20 | 7071 | 7071 | 7071 | 7071 | 7071 | 8520 | 0
2017-07-18 02:30 | 7054 | 7054 | 7054 | 7054 | 7054 | 8520 | 0
2017-07-18 02:40 | 7037 | 7037 | 7037 | 7037 | 7037 | 8520 | 0
2017-07-18 02:50 | 7022 | 7022 | 7022 | 7022 | 7022 | 8520 | 0
2017-07-18 03:00 | 7010 | 7010 | 7010 | 7010 | 7010 | 8520 | 0
2017-07-18 03:10 | 6997 | 6997 | 6997 | 6997 | 6997 | 8520 | 0
2017-07-18 03:20 | 6988 | 6988 | 6988 | 6988 | 6988 | 8520 | 0
2017-07-18 03:30 | 6982 | 6982 | 6982 | 6982 | 6982 | 8520 | 0
2017-07-18 03:40 | 6981 | 6981 | 6981 | 6981 | 6981 | 8520 | 0
2017-07-18 03:50 | 6981 | 6981 | 6981 | 6981 | 6981 | 8520 | 0
2017-07-18 04:00 | 6988 | 6988 | 6988 | 6988 | 6988 | 8520 | 0
2017-07-18 04:10 | 7002 | 7002 | 7002 | 7002 | 7002 | 8520 | 0
2017-07-18 04:20 | 7025 | 7025 | 7025 | 7025 | 7025 | 8520 | 0
2017-07-18 04:30 | 7048 | 7048 | 7048 | 7048 | 7048 | 8520 | 0
2017-07-18 04:40 | 7079 | 7079 | 7079 | 7079 | 7079 | 8520 | 0
2017-07-18 04:50 | 7112 | 7112 | 7112 | 7112 | 7112 | 8520 | 0
2017-07-18 05:00 | 7146 | 7146 | 7146 | 7146 | 7146 | 8520 | 0
2017-07-18 05:10 | 7183 | 7183 | 7183 | 7183 | 7183 | 8520 | 0
2017-07-18 05:20 | 7220 | 7220 | 7220 | 7220 | 7220 | 8520 | 0
2017-07-18 05:30 | 7260 | 7260 | 7260 | 7260 | 7260 | 8520 | 0
2017-07-18 05:40 | 7301 | 7301 | 7301 | 7301 | 7301 | 8520 | 0
2017-07-18 05:50 | 7344 | 7344 | 7344 | 7344 | 7344 | 8520 | 0
2017-07-18 06:00 | 7389 | 7389 | 7389 | 7389 | 7389 | 8520 | 0
2017-07-18 06:10 | 7435 | 7435 | 7435 | 7435 | 7435 | 8520 | 0
2017-07-18 06:20 | 7483 | 7483 | 7483 | 7483 | 7483 | 8520 | 0
2017-07-18 06:30 | 7533 | 7533 | 7533 | 7533 | 7533 | 8520 | 0
2017-07-18 06:40 | 7584 | 7584 | 7584 | 7584 | 7584 | 8520 | 0
2017-07-18 06:50 | 7636 | 7636 | 7636 | 7636 | 7636 | 8520 | 0
2017-07-18 07:00 | 7690 | 7690 | 7690 | 7690 | 7690 | 8521 | 0
2017-07-18 07:10 | 7746 | 7746 | 7746 | 7746 | 7746 | 8521 | 0
2017-07-18 07:20 | 7803 | 7803 | 7803 | 7803 | 7803 | 8521 | 0
2017-07-18 07:30 | 7861 | 7861 | 7861 | 7861 | 7861 | 8521 | 0
2017-07-18 07:40 | 7922 | 7922 | 7922 | 7922 | 7922 | 8521 | 0
2017-07-18 07:50 | 7983 | 7983 | 7983 | 7983 | 7983 | 8521 | 0
2017-07-18 08:00 | 8046 | 8046 | 8046 | 8046 | 8046 | 8521 | 0
2017-07-18 08:10 | 8111 | 8111 | 8111 | 8111 | 8111 | 8521 | 0
2017-07-18 08:20 | 8176 | 8176 | 8176 | 8176 | 8176 | 8521 | 0
2017-07-18 08:30 | 8243 | 8243 | 8243 | 8243 | 8243 | 8521 | 0
2017-07-18 08:40 | 8311 | 8311 | 8311 | 8311 | 8311 | 8521 | 0
2017-07-18 08:50 | 8380 | 8380 | 8380 | 8380 | 8380 | 8521 | 0
2017-07-18 09:00 | 8449 | 8449 | 8449 | 8449 | 8449 | 8521 | 0
2017-07-18 09:10 | 8520 | 8520 | 8520 | 8520 | 8520 | 8521 | 0
2017-07-18 09:20 | 8590 | 8590 | 8590 | 8590 | 8590 | 8521 | 0
2017-07-18 09:30 | 8661 | 8661 | 8661 | 8661 | 8661 | 8521 | 0
2017-07-18 09:40 | 8732 | 8732 | 8732 | 8732 | 8732 | 8521 | 0
2017-07-18 09:50 | 8803 | 8803 | 8803 | 8803 | 8803 | 8521 | 0
2017-07-18 10:00 | 8874 | 8874 | 8874 | 8874 | 8874 | 8522 | 0
2017-07-18 10:10 | 8945 | 8945 | 8945 | 8945 | 8945 | 8522 | 0
2017-07-18 10:20 | 9017 | 9017 | 9017 | 9017 | 9017 | 8522 | 0
2017-07-18 10:30 | 9090 | 9090 | 9090 | 9090 | 9090 | 8522 | 0
2017-07-18 10:40 | 9163 | 9163 | 9163 | 9163 | 9163 | 8522 | 0
2017-07-18 10:50 | 9236 | 9236 | 9236 | 9236 | 9236 | 8522 | 0
2017-07-18 11:00 | 9309 | 9309 | 9309 | 9309 | 9309 | 8522 | 0
2017-07-18 11:10 | 9383 | 9383 | 9383 | 9383 | 9383 | 8522 | 0
2017-07-18 11:20 | 9457 | 9457 | 9457 | 9457 | 9457 | 8522 | 0
2017-07-18 11:30 | 9531 | 9531 | 9531 | 9531 | 9531 | 8522 | 0
2017-07-18 11:40 | 9606 | 9606 | 9606 | 9606 | 9606 | 8522 | 0
2017-07-18 11:50 | 9681 | 9681 | 9681 | 9681 | 9681 | 8522 | 0
2017-07-18 12:00 | 9756 | 9756 | 9756 | 9756 | 9756 | 8522 | 0
2017-07-18 12:10 | 9831 | 9831 | 9831 | 9831 | 9831 | 8522 | 0
2017-07-18 12:20 | 9906 | 9906 | 9906 | 9906 | 9906 | 8522 | 0
2017-07-18 12:30 | 9980 | 9980 | 9980 | 9980 | 9980 | 8522 | 0
2017-07-18 12:40 | 10055 | 10055 | 10055 | 10055 | 10055 | 8522 | 0
2017-07-18 12:50 | 10129 | 10129 | 10129 | 10129 | 10129 | 8522 | 0
2017-07-18 13:00 | 10203 | 10203 | 10203 | 10203 | 10203 | 8522 | 0
2017-07-18 13:10 | 10277 | 10277 | 10277 | 10277 | 10277 | 8522 | 0
2017-07-18 13:20 | 10350 | 10350 | 10350 | 10350 | 10350 | 8522 | 0
2017-07-18 13:30 | 10422 | 10422 | 10422 | 10422 | 10422 | 8522 | 0
2017-07-18 13:40 | 10494 | 10494 | 10494 | 10494 | 10494 | 8522 | 0
2017-07-18 13:50 | 10563 | 10563 | 10563 | 10563 | 10563 | 8522 | 0
2017-07-18 14:00 | 10631 | 10631 | 10631 | 10631 | 10631 | 8522 | 0
2017-07-18 14:10 | 10695 | 10695 | 10695 | 10695 | 10695 | 8522 | 0
2017-07-18 14:20 | 10756 | 10756 | 10756 | 10756 | 10756 | 8522 | 0
2017-07-18 14:30 | 10811 | 10811 | 10811 | 10811 | 10811 | 8522 | 0
2017-07-18 14:40 | 10859 | 10859 | 10859 | 10859 | 10859 | 8522 | 0
2017-07-18 14:50 | 10900 | 10900 | 10900 | 10900 | 10900 | 8522 | 0
2017-07-18 15:00 | 10935 | 10935 | 10935 | 10935 | 10935 | 8522 | 0
2017-07-18 15:10 | 10965 | 10965 | 10965 | 10965 | 10965 | 8522 | 0
2017-07-18 15:20 | 10991 | 10991 | 10991 | 10991 | 10991 | 8522 | 0
2017-07-18 15:30 | 11010 | 11010 | 11010 | 11010 | 11010 | 8522 | 0
2017-07-18 15:40 | 11023 | 11023 | 11023 | 11023 | 11023 | 8522 | 0
2017-07-18 15:50 | 11029 | 11029 | 11029 | 11029 | 11029 | 8522 | 0
2017-07-18 16:00 | 11029 | 11029 | 11029 | 11029 | 11029 | 8522 | 0
2017-07-18 16:10 | 11021 | 11021 | 11021 | 11021 | 11021 | 8522 | 0
2017-07-18 16:20 | 11005 | 11005 | 11005 | 11005 | 11005 | 8522 | 0
2017-07-18 16:30 | 10982 | 10982 | 10982 | 10982 | 10982 | 8522 | 0
2017-07-18 16:40 | 10951 | 10951 | 10951 | 10951 | 10951 | 8522 | 0
2017-07-18 16:50 | 10912 | 10912 | 10912 | 10912 | 10912 | 8522 | 0
2017-07-18 17:00 | 10865 | 10865 | 10865 | 10865 | 10865 | 8522 | 0
2017-07-18 17:10 | 10811 | 10811 | 10811 | 10811 | 10811 | 8522 | 0
2017-07-18 17:20 | 10750 | 10750 | 10750 | 10750 | 10750 | 8522 | 0
2017-07-18 17:30 | 10682 | 10682 | 10682 | 10682 | 10682 | 8522 | 0
2017-07-18 17:40 | 10608 | 10608 | 10608 | 10608 | 10608 | 8522 | 0
2017-07-18 17:50 | 10529 | 10529 | 10529 | 10529 | 10529 | 8522 | 0
2017-07-18 18:00 | 10445 | 10445 | 10445 | 10445 | 10445 | 8522 | 0
2017-07-18 18:10 | 10357 | 10357 | 10357 | 10357 | 10357 | 8522 | 0
2017-07-18 18:20 | 10265 | 10265 | 10265 | 10265 | 10265 | 8522 | 0
2017-07-18 18:30 | 10170 | 10170 | 10170 | 10170 | 10170 | 8522 | 0
2017-07-18 18:40 | 10075 | 10075 | 10075 | 10075 | 10075 | 8522 | 0
2017-07-18 18:50 | 9981 | 9981 | 9981 | 9981 | 9981 | 8522 | 0
2017-07-18 19:00 | 9889 | 9889 | 9889 | 9889 | 9889 | 8522 | 0
2017-07-18 19:10 | 9801 | 9801 | 9801 | 9801 | 9801 | 8522 | 0
2017-07-18 19:20 | 9713 | 9713 | 9713 | 9713 | 9713 | 8522 | 0
2017-07-18 19:30 | 9624 | 9624 | 9624 | 9624 | 9624 | 8522 | 0
2017-07-18 19:40 | 9533 | 9533 | 9533 | 9533 | 9533 | 8522 | 0
2017-07-18 19:50 | 9438 | 9438 | 9438 | 9438 | 9438 | 8522 | 0
2017-07-18 20:00 | 9341 | 9341 | 9341 | 9341 | 9341 | 8522 | 0
2017-07-18 20:10 | 9240 | 9240 | 9240 | 9240 | 9240 | 8522 | 0
2017-07-18 20:20 | 9136 | 9136 | 9136 | 9136 | 9136 | 8522 | 0
2017-07-18 20:30 | 9029 | 9029 | 9029 | 9029 | 9029 | 8522 | 0
2017-07-18 20:40 | 8919 | 8919 | 8919 | 8919 | 8919 | 8522 | 0
2017-07-18 20:50 | 8806 | 8806 | 8806 | 8806 | 8806 | 8522 | 0
2017-07-18 21:00 | 8691 | 8691 | 8691 | 8691 | 8691 | 8521 | 0
2017-07-18 21:10 | 8574 | 8574 | 8574 | 8574 | 8574 | 8521 | 0
2017-07-18 21:20 | 8456 | 8456 | 8456 | 8456 | 8456 | 8521 | 0
2017-07-18 21:30 | 8337 | 8337 | 8337 | 8337 | 8337 | 8521 | 0
2017-07-18 21:40 | 8220 | 8220 | 8220 | 8220 | 8220 | 8521 | 0
2017-07-18 21:50 | 8106 | 8106 | 8106 | 8106 | 8106 | 8521 | 0
2017-07-18 22:00 | 7996 | 7996 | 7996 | 7996 | 7996 | 8521 | 0
2017-07-18 22:10 | 7892 | 7892 | 7892 | 7892 | 7892 | 8521 | 0
2017-07-18 22:20 | 7795 | 7795 | 7795 | 7795 | 7795 | 8521 | 0
2017-07-18 22:30 | 7705 | 7705 | 7705 | 7705 | 7705 | 8521 | 0
2017-07-18 22:40 | 7623 | 7623 | 7623 | 7623 | 7623 | 8521 | 0
2017-07-18 22:50 | 7546 | 7546 | 7546 | 7546 | 7546 | 8520 | 0
2017-07-18 23:00 | 7477 | 7477 | 7477 | 7477 | 7477 | 8520 | 0
2017-07-18 23:10 | 7413 | 7413 | 7413 | 7413 | 7413 | 8520 | 0
2017-07-18 23:20 | 7355 | 7355 | 7355 | 7355 | 7355 | 8520 | 0
2017-07-18 23:30 | 7303 | 7303 | 7303 | 7303 | 7303 | 8520 | 0
2017-07-18 23:40 | 7256 | 7256 | 7256 | 7256 | 7256 | 8520 | 0
2017-07-18 23:50 | 7213 | 7213 | 7213 | 7213 | 7213 | 8520 | 0
2017-07-19 | 7175 | 7175 | 7175 | 7175 | 7175 | 8520 | 0
2017-07-19 00:10 | 7141 | 7141 | 7141 | 7141 | 7141 | 8520 | 0
2017-07-19 00:20 | 7111 | 7111 | 7111 | 7111 | 7111 | 8520 | 0
2017-07-19 00:30 | 7084 | 7084 | 7084 | 7084 | 7084 | 8520 | 0
2017-07-19 00:40 | 7060 | 7060 | 7060 | 7060 | 7060 | 8520 | 0
2017-07-19 00:50 | 7039 | 7039 | 7039 | 7039 | 7039 | 8520 | 0
2017-07-19 01:00 | 7021 | 7021 | 7021 | 7021 | 7021 | 8520 | 0
2017-07-19 01:10 | 7006 | 7006 | 7006 | 7006 | 7006 | 8520 | 0
2017-07-19 01:20 | 6992 | 6992 | 6992 | 6992 | 6992 | 8520 | 0
2017-07-19 01:30 | 6981 | 6981 | 6981 | 6981 | 6981 | 8520 | 0
2017-07-19 01:40 | 6973 | 6973 | 6973 | 6973 | 6973 | 8520 | 0
2017-07-19 01:50 | 6966 | 6966 | 6966 | 6966 | 6966 | 8520 | 0
2017-07-19 02:00 | 6962 | 6962 | 6962 | 6962 | 6962 | 8520 | 0
2017-07-19 02:10 | 6959 | 6959 | 6959 | 6959 | 6959 | 8520 | 0
2017-07-19 02:20 | 6959 | 6959 | 6959 | 6959 | 6959 | 8520 | 0
2017-07-19 02:30 | 6960 | 6960 | 6960 | 6960 | 6960 | 8520 | 0
2017-07-19 02:40 | 6963 | 6963 | 6963 | 6963 | 6963 | 8520 | 0
2017-07-19 02:50 | 6969 | 6969 | 6969 | 6969 | 6969 | 8520 | 0
2017-07-19 03:00 | 6976 | 6976 | 6976 | 6976 | 6976 | 8520 | 0
2017-07-19 03:10 | 6985 | 6985 | 6985 | 6985 | 6985 | 8520 | 0
2017-07-19 03:20 | 6997 | 6997 | 6997 | 6997 | 6997 | 8520 | 0
2017-07-19 03:30 | 7025 | 7025 | 7025 | 7025 | 7025 | 8520 | 0
2017-07-19 03:40 | 7143 | 7143 | 7143 | 7143 | 7143 | 8520 | 0
2017-07-19 03:50 | 7249 | 7249 | 7249 | 7249 | 7249 | 8520 | 0
2017-07-19 04:00 | 7354 | 7354 | 7354 | 7354 | 7354 | 8520 | 0
2017-07-19 04:10 | 7457 | 7457 | 7457 | 7457 | 7457 | 8520 | 0
2017-07-19 04:20 | 7560 | 7560 | 7560 | 7560 | 7560 | 8520 | 0
2017-07-19 04:30 | 7662 | 7662 | 7662 | 7662 | 7662 | 8520 | 0
2017-07-19 04:40 | 7763 | 7763 | 7763 | 7763 | 7763 | 8521 | 0
2017-07-19 04:50 | 7864 | 7864 | 7864 | 7864 | 7864 | 8521 | 0
2017-07-19 05:00 | 7962 | 7962 | 7962 | 7962 | 7962 | 8521 | 0
2017-07-19 05:10 | 8059 | 8059 | 8059 | 8059 | 8059 | 8521 | 0
2017-07-19 05:20 | 8153 | 8153 | 8153 | 8153 | 8153 | 8521 | 0
2017-07-19 05:30 | 8245 | 8245 | 8245 | 8245 | 8245 | 8521 | 0
2017-07-19 05:40 | 8337 | 8337 | 8337 | 8337 | 8337 | 8521 | 0
2017-07-19 05:50 | 8426 | 8426 | 8426 | 8426 | 8426 | 8521 | 0
2017-07-19 06:00 | 8513 | 8513 | 8513 | 8513 | 8513 | 8521 | 0
2017-07-19 06:10 | 8598 | 8598 | 8598 | 8598 | 8598 | 8521 | 0
2017-07-19 06:20 | 8681 | 8681 | 8681 | 8681 | 8681 | 8521 | 0
2017-07-19 06:30 | 8762 | 8762 | 8762 | 8762 | 8762 | 8521 | 0
2017-07-19 06:40 | 8840 | 8840 | 8840 | 8840 | 8840 | 8521 | 0
2017-07-19 06:50 | 8918 | 8918 | 8918 | 8918 | 8918 | 8522 | 0
2017-07-19 07:00 | 8994 | 8994 | 8994 | 8994 | 8994 | 8522 | 0
2017-07-19 07:10 | 9069 | 9069 | 9069 | 9069 | 9069 | 8522 | 0
2017-07-19 07:20 | 9143 | 9143 | 9143 | 9143 | 9143 | 8522 | 0
2017-07-19 07:30 | 9216 | 9216 | 9216 | 9216 | 9216 | 8522 | 0
2017-07-19 07:40 | 9290 | 9290 | 9290 | 9290 | 9290 | 8522 | 0
2017-07-19 07:50 | 9366 | 9366 | 9366 | 9366 | 9366 | 8522 | 0
2017-07-19 08:00 | 9441 | 9441 | 9441 | 9441 | 9441 | 8522 | 0
2017-07-19 08:10 | 9516 | 9516 | 9516 | 9516 | 9516 | 8522 | 0
2017-07-19 08:20 | 9589 | 9589 | 9589 | 9589 | 9589 | 8522 | 0
2017-07-19 08:30 | 9661 | 9661 | 9661 | 9661 | 9661 | 8522 | 0
2017-07-19 08:40 | 9731 | 9731 | 9731 | 9731 | 9731 | 8522 | 0
2017-07-19 08:50 | 9799 | 9799 | 9799 | 9799 | 9799 | 8522 | 0
2017-07-19 09:00 | 9868 | 9868 | 9868 | 9868 | 9868 | 8522 | 0
2017-07-19 09:10 | 9937 | 9937 | 9937 | 9937 | 9937 | 8522 | 0
2017-07-19 09:20 | 10006 | 10006 | 10006 | 10006 | 10006 | 8522 | 0
2017-07-19 09:30 | 10074 | 10074 | 10074 | 10074 | 10074 | 8522 | 0
2017-07-19 09:40 | 10141 | 10141 | 10141 | 10141 | 10141 | 8522 | 0
2017-07-19 09:50 | 10206 | 10206 | 10206 | 10206 | 10206 | 8522 | 0
2017-07-19 10:00 | 10271 | 10271 | 10271 | 10271 | 10271 | 8522 | 0
2017-07-19 10:10 | 10335 | 10335 | 10335 | 10335 | 10335 | 8522 | 0
2017-07-19 10:20 | 10398 | 10398 | 10398 | 10398 | 10398 | 8522 | 0
2017-07-19 10:30 | 10460 | 10460 | 10460 | 10460 | 10460 | 8522 | 0
2017-07-19 10:40 | 10520 | 10520 | 10520 | 10520 | 10520 | 8522 | 0
2017-07-19 10:50 | 10579 | 10579 | 10579 | 10579 | 10579 | 8522 | 0
2017-07-19 11:00 | 10638 | 10638 | 10638 | 10638 | 10638 | 8522 | 0
2017-07-19 11:10 | 10699 | 10699 | 10699 | 10699 | 10699 | 8522 | 0
2017-07-19 11:20 | 10760 | 10760 | 10760 | 10760 | 10760 | 8522 | 0
2017-07-19 11:30 | 10821 | 10821 | 10821 | 10821 | 10821 | 8522 | 0
2017-07-19 11:40 | 10881 | 10881 | 10881 | 10881 | 10881 | 8522 | 0
2017-07-19 11:50 | 10941 | 10941 | 10941 | 10941 | 10941 | 8522 | 0
2017-07-19 12:00 | 11001 | 11001 | 11001 | 11001 | 11001 | 8522 | 0
2017-07-19 12:10 | 11077 | 11077 | 11077 | 11077 | 11077 | 8522 | 0
2017-07-19 12:20 | 11171 | 11171 | 11171 | 11171 | 11171 | 8522 | 0
2017-07-19 12:30 | 11221 | 11221 | 11221 | 11221 | 11221 | 8522 | 0
2017-07-19 12:40 | 11271 | 11271 | 11271 | 11271 | 11271 | 8522 | 0
2017-07-19 12:50 | 11324 | 11324 | 11324 | 11324 | 11324 | 8522 | 0
2017-07-19 13:00 | 11378 | 11378 | 11378 | 11378 | 11378 | 8522 | 0
2017-07-19 13:10 | 11429 | 11429 | 11429 | 11429 | 11429 | 8522 | 0
2017-07-19 13:20 | 11477 | 11477 | 11477 | 11477 | 11477 | 8522 | 0
2017-07-19 13:30 | 11499 | 11499 | 11499 | 11499 | 11499 | 8522 | 0
2017-07-19 13:40 | 11431 | 11431 | 11431 | 11431 | 11431 | 8522 | 0
2017-07-19 13:50 | 11271 | 11271 | 11271 | 11271 | 11271 | 8522 | 0
2017-07-19 14:00 | 11042 | 11042 | 11042 | 11042 | 11042 | 8522 | 0
2017-07-19 14:10 | 10803 | 10803 | 10803 | 10803 | 10803 | 8522 | 0
2017-07-19 14:20 | 10605 | 10605 | 10605 | 10605 | 10605 | 8522 | 0
2017-07-19 14:30 | 10464 | 10464 | 10464 | 10464 | 10464 | 8522 | 0
2017-07-19 14:40 | 10374 | 10374 | 10374 | 10374 | 10374 | 8522 | 0
2017-07-19 14:50 | 10355 | 10355 | 10355 | 10355 | 10355 | 8522 | 0
2017-07-19 15:00 | 10289 | 10289 | 10289 | 10289 | 10289 | 8522 | 0
2017-07-19 15:10 | 10159 | 10159 | 10159 | 10159 | 10159 | 8522 | 0
2017-07-19 15:20 | 10040 | 10040 | 10040 | 10040 | 10040 | 8522 | 0
2017-07-19 15:30 | 9966 | 9966 | 9966 | 9966 | 9966 | 8522 | 0
2017-07-19 15:40 | 9921 | 9921 | 9921 | 9921 | 9921 | 8522 | 0
2017-07-19 15:50 | 9875 | 9875 | 9875 | 9875 | 9875 | 8522 | 0
2017-07-19 16:00 | 9785 | 9785 | 9785 | 9785 | 9785 | 8522 | 0
2017-07-19 16:10 | 9643 | 9643 | 9643 | 9643 | 9643 | 8522 | 0
2017-07-19 16:20 | 9471 | 9471 | 9471 | 9471 | 9471 | 8522 | 0
2017-07-19 16:30 | 9309 | 9309 | 9309 | 9309 | 9309 | 8522 | 0
2017-07-19 16:40 | 9261 | 9261 | 9261 | 9261 | 9261 | 8522 | 0
2017-07-19 16:50 | 9239 | 9239 | 9239 | 9239 | 9239 | 8522 | 0
2017-07-19 17:00 | 9212 | 9212 | 9212 | 9212 | 9212 | 8522 | 0
2017-07-19 17:10 | 9137 | 9137 | 9137 | 9137 | 9137 | 8522 | 0
2017-07-19 17:20 | 9266 | 9266 | 9266 | 9266 | 9266 | 8522 | 0
2017-07-19 17:30 | 9413 | 9413 | 9413 | 9413 | 9413 | 8522 | 0
2017-07-19 17:40 | 9476 | 9476 | 9476 | 9476 | 9476 | 8522 | 0
2017-07-19 17:50 | 9494 | 9494 | 9494 | 9494 | 9494 | 8522 | 0
2017-07-19 18:00 | 9513 | 9513 | 9513 | 9513 | 9513 | 8522 | 0
2017-07-19 18:10 | 9527 | 9527 | 9527 | 9527 | 9527 | 8522 | 0
2017-07-19 18:20 | 9540 | 9540 | 9540 | 9540 | 9540 | 8522 | 0
2017-07-19 18:30 | 9569 | 9569 | 9569 | 9569 | 9569 | 8522 | 0
2017-07-19 18:40 | 9669 | 9669 | 9669 | 9669 | 9669 | 8522 | 0
2017-07-19 18:50 | 9891 | 9891 | 9891 | 9891 | 9891 | 8522 | 0
2017-07-19 19:00 | 10213 | 10213 | 10213 | 10213 | 10213 | 8522 | 0
2017-07-19 19:10 | 10601 | 10601 | 10601 | 10601 | 10601 | 8522 | 0
2017-07-19 19:20 | 11036 | 11036 | 11036 | 11036 | 11036 | 8522 | 0
2017-07-19 19:30 | 11508 | 11508 | 11508 | 11508 | 11508 | 8522 | 0
2017-07-19 19:40 | 12001 | 12001 | 12001 | 12001 | 12001 | 8522 | 0
2017-07-19 19:50 | 12508 | 12508 | 12508 | 12508 | 12508 | 8522 | 0
2017-07-19 20:00 | 13026 | 13026 | 13026 | 13026 | 13026 | 8522 | 0
2017-07-19 20:10 | 13556 | 13556 | 13556 | 13556 | 13556 | 8522 | 0
2017-07-19 20:20 | 14089 | 14089 | 14089 | 14089 | 14089 | 8522 | 0
2017-07-19 20:30 | 14621 | 14621 | 14621 | 14621 | 14621 | 8522 | 0
2017-07-19 20:40 | 15149 | 15149 | 15149 | 15149 | 15149 | 8522 | 0
2017-07-19 20:50 | 15668 | 15668 | 15668 | 15668 | 15668 | 8522 | 0
2017-07-19 21:00 | 16124 | 16124 | 16124 | 16124 | 16124 | 8522 | 0
2017-07-19 21:10 | 16509 | 16509 | 16509 | 16509 | 16509 | 8522 | 0
2017-07-19 21:20 | 16691 | 16691 | 16691 | 16691 | 16691 | 8522 | 0
2017-07-19 21:30 | 16236 | 16236 | 16236 | 16236 | 16236 | 8522 | 0
2017-07-19 21:40 | 15510 | 15510 | 15510 | 15510 | 15510 | 8522 | 0
2017-07-19 21:50 | 15080 | 15080 | 15080 | 15080 | 15080 | 8522 | 0
2017-07-19 22:00 | 14749 | 14749 | 14749 | 14749 | 14749 | 8522 | 0
2017-07-19 22:10 | 14524 | 14524 | 14524 | 14524 | 14524 | 8522 | 0
2017-07-19 22:20 | 14325 | 14325 | 14325 | 14325 | 14325 | 8522 | 0
2017-07-19 22:30 | 14139 | 14139 | 14139 | 14139 | 14139 | 8522 | 0
2017-07-19 22:40 | 13978 | 13978 | 13978 | 13978 | 13978 | 8522 | 0
2017-07-19 22:50 | 13826 | 13826 | 13826 | 13826 | 13826 | 8522 | 0
2017-07-19 23:00 | 13614 | 13614 | 13614 | 13614 | 13614 | 8522 | 0
2017-07-19 23:10 | 13400 | 13400 | 13400 | 13400 | 13400 | 8522 | 0
2017-07-19 23:20 | 13202 | 13202 | 13202 | 13202 | 13202 | 8523 | 0
2017-07-19 23:30 | 13019 | 13019 | 13019 | 13019 | 13019 | 8522 | 0
2017-07-19 23:40 | 12860 | 12860 | 12860 | 12860 | 12860 | 8522 | 0
2017-07-19 23:50 | 12716 | 12716 | 12716 | 12716 | 12716 | 8522 | 0
2017-07-20 | 12581 | 12581 | 12581 | 12581 | 12581 | 8523 | 0
2017-07-20 00:10 | 12453 | 12453 | 12453 | 12453 | 12453 | 8523 | 0
2017-07-20 00:20 | 12331 | 12331 | 12331 | 12331 | 12331 | 8522 | 0
2017-07-20 00:30 | 12226 | 12226 | 12226 | 12226 | 12226 | 8523 | 0
2017-07-20 00:40 | 12125 | 12125 | 12125 | 12125 | 12125 | 8523 | 0
2017-07-20 00:50 | 12038 | 12038 | 12038 | 12038 | 12038 | 8523 | 0
2017-07-20 01:00 | 11953 | 11953 | 11953 | 11953 | 11953 | 8523 | 0
2017-07-20 01:10 | 11871 | 11871 | 11871 | 11871 | 11871 | 8522 | 0
2017-07-20 01:20 | 11795 | 11795 | 11795 | 11795 | 11795 | 8522 | 0
2017-07-20 01:30 | 11723 | 11723 | 11723 | 11723 | 11723 | 8522 | 0
2017-07-20 01:40 | 11654 | 11654 | 11654 | 11654 | 11654 | 8522 | 0
2017-07-20 01:50 | 11620 | 11620 | 11620 | 11620 | 11620 | 8522 | 0
2017-07-20 02:00 | 11586 | 11586 | 11586 | 11586 | 11586 | 8522 | 0
2017-07-20 02:10 | 11670 | 11670 | 11670 | 11670 | 11670 | 8522 | 0
2017-07-20 02:20 | 11583 | 11583 | 11583 | 11583 | 11583 | 8522 | 0
2017-07-20 02:30 | 11529 | 11529 | 11529 | 11529 | 11529 | 8522 | 0
2017-07-20 02:40 | 11477 | 11477 | 11477 | 11477 | 11477 | 8522 | 0
2017-07-20 02:50 | 11429 | 11429 | 11429 | 11429 | 11429 | 8522 | 0
2017-07-20 03:00 | 11375 | 11375 | 11375 | 11375 | 11375 | 8522 | 0
2017-07-20 03:10 | 11323 | 11323 | 11323 | 11323 | 11323 | 8522 | 0
2017-07-20 03:20 | 11274 | 11274 | 11274 | 11274 | 11274 | 8522 | 0
2017-07-20 03:30 | 11226 | 11226 | 11226 | 11226 | 11226 | 8522 | 0
2017-07-20 03:40 | 11182 | 11182 | 11182 | 11182 | 11182 | 8522 | 0
2017-07-20 03:50 | 11140 | 11140 | 11140 | 11140 | 11140 | 8522 | 0
2017-07-20 04:00 | 11100 | 11100 | 11100 | 11100 | 11100 | 8522 | 0
2017-07-20 04:10 | 11062 | 11062 | 11062 | 11062 | 11062 | 8522 | 0
2017-07-20 04:20 | 10960 | 10960 | 10960 | 10960 | 10960 | 8522 | 0
2017-07-20 04:30 | 10845 | 10845 | 10845 | 10845 | 10845 | 8522 | 0
2017-07-20 04:40 | 10779 | 10779 | 10779 | 10779 | 10779 | 8522 | 0
2017-07-20 04:50 | 10720 | 10720 | 10720 | 10720 | 10720 | 8522 | 0
2017-07-20 05:00 | 10658 | 10658 | 10658 | 10658 | 10658 | 8522 | 0
2017-07-20 05:10 | 10519 | 10519 | 10519 | 10519 | 10519 | 8522 | 0
2017-07-20 05:20 | 10438 | 10438 | 10438 | 10438 | 10438 | 8522 | 0
2017-07-20 05:30 | 10406 | 10406 | 10406 | 10406 | 10406 | 8522 | 0
2017-07-20 05:40 | 10388 | 10388 | 10388 | 10388 | 10388 | 8522 | 0
2017-07-20 05:50 | 10375 | 10375 | 10375 | 10375 | 10375 | 8522 | 0
2017-07-20 06:00 | 10365 | 10365 | 10365 | 10365 | 10365 | 8522 | 0
2017-07-20 06:10 | 10359 | 10359 | 10359 | 10359 | 10359 | 8522 | 0
2017-07-20 06:20 | 10356 | 10356 | 10356 | 10356 | 10356 | 8522 | 0
2017-07-20 06:30 | 10355 | 10355 | 10355 | 10355 | 10355 | 8522 | 0
2017-07-20 06:40 | 10352 | 10352 | 10352 | 10352 | 10352 | 8522 | 0
2017-07-20 06:50 | 10350 | 10350 | 10350 | 10350 | 10350 | 8522 | 0
2017-07-20 07:00 | 10347 | 10347 | 10347 | 10347 | 10347 | 8522 | 0
2017-07-20 07:10 | 10342 | 10342 | 10342 | 10342 | 10342 | 8522 | 0
2017-07-20 07:20 | 10337 | 10337 | 10337 | 10337 | 10337 | 8522 | 0
2017-07-20 07:30 | 10331 | 10331 | 10331 | 10331 | 10331 | 8522 | 0
2017-07-20 07:40 | 10325 | 10325 | 10325 | 10325 | 10325 | 8522 | 0
2017-07-20 07:50 | 10319 | 10319 | 10319 | 10319 | 10319 | 8522 | 0
2017-07-20 08:00 | 10313 | 10313 | 10313 | 10313 | 10313 | 8522 | 0
2017-07-20 08:10 | 10307 | 10307 | 10307 | 10307 | 10307 | 8522 | 0
2017-07-20 08:20 | 10300 | 10300 | 10300 | 10300 | 10300 | 8522 | 0
2017-07-20 08:30 | 10294 | 10294 | 10294 | 10294 | 10294 | 8522 | 0
2017-07-20 08:40 | 10288 | 10288 | 10288 | 10288 | 10288 | 8522 | 0
2017-07-20 08:50 | 10282 | 10282 | 10282 | 10282 | 10282 | 8522 | 0
2017-07-20 09:00 | 10276 | 10276 | 10276 | 10276 | 10276 | 8522 | 0
2017-07-20 09:10 | 10270 | 10270 | 10270 | 10270 | 10270 | 8522 | 0
2017-07-20 09:20 | 10264 | 10264 | 10264 | 10264 | 10264 | 8522 | 0
2017-07-20 09:30 | 10259 | 10259 | 10259 | 10259 | 10259 | 8522 | 0
2017-07-20 09:40 | 10253 | 10253 | 10253 | 10253 | 10253 | 8522 | 0
2017-07-20 09:50 | 10248 | 10248 | 10248 | 10248 | 10248 | 8522 | 0
2017-07-20 10:00 | 10242 | 10242 | 10242 | 10242 | 10242 | 8522 | 0
2017-07-20 10:10 | 10237 | 10237 | 10237 | 10237 | 10237 | 8522 | 0
2017-07-20 10:20 | 10232 | 10232 | 10232 | 10232 | 10232 | 8522 | 0
2017-07-20 10:30 | 10227 | 10227 | 10227 | 10227 | 10227 | 8522 | 0
2017-07-20 10:40 | 10222 | 10222 | 10222 | 10222 | 10222 | 8522 | 0
2017-07-20 10:50 | 10217 | 10217 | 10217 | 10217 | 10217 | 8522 | 0
2017-07-20 11:00 | 10212 | 10212 | 10212 | 10212 | 10212 | 8522 | 0
2017-07-20 11:10 | 10207 | 10207 | 10207 | 10207 | 10207 | 8522 | 0
2017-07-20 11:20 | 10202 | 10202 | 10202 | 10202 | 10202 | 8522 | 0
2017-07-20 11:30 | 10197 | 10197 | 10197 | 10197 | 10197 | 8522 | 0
2017-07-20 11:40 | 10193 | 10193 | 10193 | 10193 | 10193 | 8522 | 0
2017-07-20 11:50 | 10188 | 10188 | 10188 | 10188 | 10188 | 8522 | 0
2017-07-20 12:00 | 10184 | 10184 | 10184 | 10184 | 10184 | 8522 | 0
2017-07-20 12:10 | 10180 | 10180 | 10180 | 10180 | 10180 | 8522 | 0
2017-07-20 12:20 | 10176 | 10176 | 10176 | 10176 | 10176 | 8522 | 0
2017-07-20 12:30 | 10173 | 10173 | 10173 | 10173 | 10173 | 8522 | 0
2017-07-20 12:40 | 10169 | 10169 | 10169 | 10169 | 10169 | 8522 | 0
2017-07-20 12:50 | 10166 | 10166 | 10166 | 10166 | 10166 | 8522 | 0
2017-07-20 13:00 | 10164 | 10164 | 10164 | 10164 | 10164 | 8522 | 0
2017-07-20 13:10 | 10161 | 10161 | 10161 | 10161 | 10161 | 8522 | 0
2017-07-20 13:20 | 10158 | 10158 | 10158 | 10158 | 10158 | 8522 | 0
2017-07-20 13:30 | 10156 | 10156 | 10156 | 10156 | 10156 | 8522 | 0
2017-07-20 13:40 | 10154 | 10154 | 10154 | 10154 | 10154 | 8522 | 0
2017-07-20 13:50 | 10151 | 10151 | 10151 | 10151 | 10151 | 8522 | 0
2017-07-20 14:00 | 10149 | 10149 | 10149 | 10149 | 10149 | 8522 | 0
2017-07-20 14:10 | 10145 | 10145 | 10145 | 10145 | 10145 | 8522 | 0
2017-07-20 14:20 | 10140 | 10140 | 10140 | 10140 | 10140 | 8522 | 0
2017-07-20 14:30 | 10134 | 10134 | 10134 | 10134 | 10134 | 8522 | 0
2017-07-20 14:40 | 10127 | 10127 | 10127 | 10127 | 10127 | 8522 | 0
2017-07-20 14:50 | 10119 | 10119 | 10119 | 10119 | 10119 | 8522 | 0
2017-07-20 15:00 | 10130 | 10130 | 10130 | 10130 | 10130 | 8522 | 0
2017-07-20 15:10 | 10637 | 10637 | 10637 | 10637 | 10637 | 8522 | 0
2017-07-20 15:20 | 11065 | 11065 | 11065 | 11065 | 11065 | 8522 | 0
2017-07-20 15:30 | 11537 | 11537 | 11537 | 11537 | 11537 | 8522 | 0
2017-07-20 15:40 | 11661 | 11661 | 11661 | 11661 | 11661 | 8522 | 0
2017-07-20 15:50 | 11531 | 11531 | 11531 | 11531 | 11531 | 8522 | 0
2017-07-20 16:00 | 11417 | 11417 | 11417 | 11417 | 11417 | 8522 | 0
2017-07-20 16:10 | 11311 | 11311 | 11311 | 11311 | 11311 | 8522 | 0
2017-07-20 16:20 | 11232 | 11232 | 11232 | 11232 | 11232 | 8522 | 0
2017-07-20 16:30 | 11144 | 11144 | 11144 | 11144 | 11144 | 8522 | 0
2017-07-20 16:40 | 10878 | 10878 | 10878 | 10878 | 10878 | 8522 | 0
2017-07-20 16:50 | 10783 | 10783 | 10783 | 10783 | 10783 | 8522 | 0
2017-07-20 17:00 | 10689 | 10689 | 10689 | 10689 | 10689 | 8522 | 0
2017-07-20 17:10 | 10646 | 10646 | 10646 | 10646 | 10646 | 8522 | 0
2017-07-20 17:20 | 10637 | 10637 | 10637 | 10637 | 10637 | 8522 | 0
2017-07-20 17:30 | 10637 | 10637 | 10637 | 10637 | 10637 | 8522 | 0
2017-07-20 17:40 | 10639 | 10639 | 10639 | 10639 | 10639 | 8522 | 0
2017-07-20 17:50 | 10642 | 10642 | 10642 | 10642 | 10642 | 8522 | 0
2017-07-20 18:00 | 10644 | 10644 | 10644 | 10644 | 10644 | 8522 | 0
2017-07-20 18:10 | 10647 | 10647 | 10647 | 10647 | 10647 | 8522 | 0
2017-07-20 18:20 | 10648 | 10648 | 10648 | 10648 | 10648 | 8522 | 0
2017-07-20 18:30 | 10649 | 10649 | 10649 | 10649 | 10649 | 8522 | 0
2017-07-20 18:40 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-20 18:50 | 10651 | 10651 | 10651 | 10651 | 10651 | 8522 | 0
2017-07-20 19:00 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 19:10 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 19:20 | 10653 | 10653 | 10653 | 10653 | 10653 | 8522 | 0
2017-07-20 19:30 | 10653 | 10653 | 10653 | 10653 | 10653 | 8522 | 0
2017-07-20 19:40 | 10654 | 10654 | 10654 | 10654 | 10654 | 8522 | 0
2017-07-20 19:50 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 20:00 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 20:10 | 10653 | 10653 | 10653 | 10653 | 10653 | 8522 | 0
2017-07-20 20:20 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 20:30 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-20 20:40 | 10648 | 10648 | 10648 | 10648 | 10648 | 8522 | 0
2017-07-20 20:50 | 10645 | 10645 | 10645 | 10645 | 10645 | 8522 | 0
2017-07-20 21:00 | 10646 | 10646 | 10646 | 10646 | 10646 | 8522 | 0
2017-07-20 21:10 | 10644 | 10644 | 10644 | 10644 | 10644 | 8522 | 0
2017-07-20 21:20 | 10644 | 10644 | 10644 | 10644 | 10644 | 8522 | 0
2017-07-20 21:30 | 10641 | 10641 | 10641 | 10641 | 10641 | 8522 | 0
2017-07-20 21:40 | 10723 | 10723 | 10723 | 10723 | 10723 | 8522 | 0
2017-07-20 21:50 | 10764 | 10764 | 10764 | 10764 | 10764 | 8522 | 0
2017-07-20 22:00 | 10781 | 10781 | 10781 | 10781 | 10781 | 8522 | 0
2017-07-20 22:10 | 10789 | 10789 | 10789 | 10789 | 10789 | 8522 | 0
2017-07-20 22:20 | 10801 | 10801 | 10801 | 10801 | 10801 | 8522 | 0
2017-07-20 22:30 | 10811 | 10811 | 10811 | 10811 | 10811 | 8522 | 0
2017-07-20 22:40 | 10806 | 10806 | 10806 | 10806 | 10806 | 8522 | 0
2017-07-20 22:50 | 10797 | 10797 | 10797 | 10797 | 10797 | 8522 | 0
2017-07-20 23:00 | 10782 | 10782 | 10782 | 10782 | 10782 | 8522 | 0
2017-07-20 23:10 | 10758 | 10758 | 10758 | 10758 | 10758 | 8522 | 0
2017-07-20 23:20 | 10739 | 10739 | 10739 | 10739 | 10739 | 8522 | 0
2017-07-20 23:30 | 10721 | 10721 | 10721 | 10721 | 10721 | 8522 | 0
2017-07-20 23:40 | 10701 | 10701 | 10701 | 10701 | 10701 | 8522 | 0
2017-07-20 23:50 | 10679 | 10679 | 10679 | 10679 | 10679 | 8522 | 0
2017-07-21 | 10656 | 10656 | 10656 | 10656 | 10656 | 8522 | 0
2017-07-21 00:10 | 10632 | 10632 | 10632 | 10632 | 10632 | 8522 | 0
2017-07-21 00:20 | 10555 | 10555 | 10555 | 10555 | 10555 | 8522 | 0
2017-07-21 00:30 | 10300 | 10300 | 10300 | 10300 | 10300 | 8522 | 0
2017-07-21 00:40 | 10117 | 10117 | 10117 | 10117 | 10117 | 8522 | 0
2017-07-21 00:50 | 10096 | 10096 | 10096 | 10096 | 10096 | 8522 | 0
2017-07-21 01:00 | 10095 | 10095 | 10095 | 10095 | 10095 | 8522 | 0
2017-07-21 01:10 | 10103 | 10103 | 10103 | 10103 | 10103 | 8522 | 0
2017-07-21 01:20 | 10106 | 10106 | 10106 | 10106 | 10106 | 8522 | 0
2017-07-21 01:30 | 10097 | 10097 | 10097 | 10097 | 10097 | 8522 | 0
2017-07-21 01:40 | 10082 | 10082 | 10082 | 10082 | 10082 | 8522 | 0
2017-07-21 01:50 | 10060 | 10060 | 10060 | 10060 | 10060 | 8522 | 0
2017-07-21 02:00 | 10047 | 10047 | 10047 | 10047 | 10047 | 8522 | 0
2017-07-21 02:10 | 10035 | 10035 | 10035 | 10035 | 10035 | 8522 | 0
2017-07-21 02:20 | 10025 | 10025 | 10025 | 10025 | 10025 | 8522 | 0
2017-07-21 02:30 | 10014 | 10014 | 10014 | 10014 | 10014 | 8522 | 0
2017-07-21 02:40 | 10004 | 10004 | 10004 | 10004 | 10004 | 8522 | 0
2017-07-21 02:50 | 9994 | 9994 | 9994 | 9994 | 9994 | 8522 | 0
2017-07-21 03:00 | 9982 | 9982 | 9982 | 9982 | 9982 | 8522 | 0
2017-07-21 03:10 | 9970 | 9970 | 9970 | 9970 | 9970 | 8522 | 0
2017-07-21 03:20 | 9963 | 9963 | 9963 | 9963 | 9963 | 8522 | 0
2017-07-21 03:30 | 9959 | 9959 | 9959 | 9959 | 9959 | 8522 | 0
2017-07-21 03:40 | 9956 | 9956 | 9956 | 9956 | 9956 | 8522 | 0
2017-07-21 03:50 | 9959 | 9959 | 9959 | 9959 | 9959 | 8522 | 0
2017-07-21 04:00 | 9962 | 9962 | 9962 | 9962 | 9962 | 8522 | 0
2017-07-21 04:10 | 9963 | 9963 | 9963 | 9963 | 9963 | 8522 | 0
2017-07-21 04:20 | 9953 | 9953 | 9953 | 9953 | 9953 | 8522 | 0
2017-07-21 04:30 | 9942 | 9942 | 9942 | 9942 | 9942 | 8522 | 0
2017-07-21 04:40 | 9933 | 9933 | 9933 | 9933 | 9933 | 8522 | 0
2017-07-21 04:50 | 9913 | 9913 | 9913 | 9913 | 9913 | 8522 | 0
2017-07-21 05:00 | 9858 | 9858 | 9858 | 9858 | 9858 | 8522 | 0
2017-07-21 05:10 | 9807 | 9807 | 9807 | 9807 | 9807 | 8522 | 0
2017-07-21 05:20 | 9812 | 9812 | 9812 | 9812 | 9812 | 8522 | 0
2017-07-21 05:30 | 9853 | 9853 | 9853 | 9853 | 9853 | 8522 | 0
2017-07-21 05:40 | 9895 | 9895 | 9895 | 9895 | 9895 | 8522 | 0
2017-07-21 05:50 | 9938 | 9938 | 9938 | 9938 | 9938 | 8522 | 0
2017-07-21 06:00 | 9982 | 9982 | 9982 | 9982 | 9982 | 8522 | 0
2017-07-21 06:10 | 10028 | 10028 | 10028 | 10028 | 10028 | 8522 | 0
2017-07-21 06:20 | 10076 | 10076 | 10076 | 10076 | 10076 | 8522 | 0
2017-07-21 06:30 | 10125 | 10125 | 10125 | 10125 | 10125 | 8522 | 0
2017-07-21 06:40 | 10169 | 10169 | 10169 | 10169 | 10169 | 8522 | 0
2017-07-21 06:50 | 10204 | 10204 | 10204 | 10204 | 10204 | 8522 | 0
2017-07-21 07:00 | 10240 | 10240 | 10240 | 10240 | 10240 | 8522 | 0
2017-07-21 07:10 | 10268 | 10268 | 10268 | 10268 | 10268 | 8522 | 0
2017-07-21 07:20 | 10301 | 10301 | 10301 | 10301 | 10301 | 8522 | 0
2017-07-21 07:30 | 10338 | 10338 | 10338 | 10338 | 10338 | 8522 | 0
2017-07-21 07:40 | 10373 | 10373 | 10373 | 10373 | 10373 | 8522 | 0
2017-07-21 07:50 | 10406 | 10406 | 10406 | 10406 | 10406 | 8522 | 0
2017-07-21 08:00 | 10436 | 10436 | 10436 | 10436 | 10436 | 8522 | 0
2017-07-21 08:10 | 10467 | 10467 | 10467 | 10467 | 10467 | 8522 | 0
2017-07-21 08:20 | 10496 | 10496 | 10496 | 10496 | 10496 | 8522 | 0
2017-07-21 08:30 | 10528 | 10528 | 10528 | 10528 | 10528 | 8522 | 0
2017-07-21 08:40 | 10559 | 10559 | 10559 | 10559 | 10559 | 8522 | 0
2017-07-21 08:50 | 10579 | 10579 | 10579 | 10579 | 10579 | 8522 | 0
2017-07-21 09:00 | 10611 | 10611 | 10611 | 10611 | 10611 | 8522 | 0
2017-07-21 09:10 | 10645 | 10645 | 10645 | 10645 | 10645 | 8522 | 0
2017-07-21 09:20 | 10681 | 10681 | 10681 | 10681 | 10681 | 8522 | 0
2017-07-21 09:30 | 10718 | 10718 | 10718 | 10718 | 10718 | 8522 | 0
2017-07-21 09:40 | 10751 | 10751 | 10751 | 10751 | 10751 | 8522 | 0
2017-07-21 09:50 | 10783 | 10783 | 10783 | 10783 | 10783 | 8522 | 0
2017-07-21 10:00 | 10813 | 10813 | 10813 | 10813 | 10813 | 8522 | 0
2017-07-21 10:10 | 10841 | 10841 | 10841 | 10841 | 10841 | 8522 | 0
2017-07-21 10:20 | 10869 | 10869 | 10869 | 10869 | 10869 | 8522 | 0
2017-07-21 10:30 | 10899 | 10899 | 10899 | 10899 | 10899 | 8522 | 0
2017-07-21 10:40 | 10929 | 10929 | 10929 | 10929 | 10929 | 8522 | 0
2017-07-21 10:50 | 10960 | 10960 | 10960 | 10960 | 10960 | 8522 | 0
2017-07-21 11:00 | 10987 | 10987 | 10987 | 10987 | 10987 | 8522 | 0
2017-07-21 11:10 | 11014 | 11014 | 11014 | 11014 | 11014 | 8522 | 0
2017-07-21 11:20 | 11040 | 11040 | 11040 | 11040 | 11040 | 8522 | 0
2017-07-21 11:30 | 11067 | 11067 | 11067 | 11067 | 11067 | 8522 | 0
2017-07-21 11:40 | 11091 | 11091 | 11091 | 11091 | 11091 | 8522 | 0
2017-07-21 11:50 | 11117 | 11117 | 11117 | 11117 | 11117 | 8522 | 0
2017-07-21 12:00 | 11142 | 11142 | 11142 | 11142 | 11142 | 8522 | 0
2017-07-21 12:10 | 11167 | 11167 | 11167 | 11167 | 11167 | 8522 | 0
2017-07-21 12:20 | 11190 | 11190 | 11190 | 11190 | 11190 | 8522 | 0
2017-07-21 12:30 | 11214 | 11214 | 11214 | 11214 | 11214 | 8522 | 0
2017-07-21 12:40 | 11238 | 11238 | 11238 | 11238 | 11238 | 8522 | 0
2017-07-21 12:50 | 11262 | 11262 | 11262 | 11262 | 11262 | 8522 | 0
2017-07-21 13:00 | 11286 | 11286 | 11286 | 11286 | 11286 | 8522 | 0
2017-07-21 13:10 | 11311 | 11311 | 11311 | 11311 | 11311 | 8522 | 0
2017-07-21 13:20 | 11336 | 11336 | 11336 | 11336 | 11336 | 8522 | 0
2017-07-21 13:30 | 11361 | 11361 | 11361 | 11361 | 11361 | 8522 | 0
2017-07-21 13:40 | 11380 | 11380 | 11380 | 11380 | 11380 | 8522 | 0
2017-07-21 13:50 | 11397 | 11397 | 11397 | 11397 | 11397 | 8522 | 0
2017-07-21 14:00 | 11404 | 11404 | 11404 | 11404 | 11404 | 8522 | 0
2017-07-21 14:10 | 11415 | 11415 | 11415 | 11415 | 11415 | 8522 | 0
2017-07-21 14:20 | 11421 | 11421 | 11421 | 11421 | 11421 | 8522 | 0
2017-07-21 14:30 | 11400 | 11400 | 11400 | 11400 | 11400 | 8522 | 0
2017-07-21 14:40 | 11402 | 11402 | 11402 | 11402 | 11402 | 8522 | 0
2017-07-21 14:50 | 11407 | 11407 | 11407 | 11407 | 11407 | 8522 | 0
2017-07-21 15:00 | 11397 | 11397 | 11397 | 11397 | 11397 | 8522 | 0
2017-07-21 15:10 | 11278 | 11278 | 11278 | 11278 | 11278 | 8522 | 0
2017-07-21 15:20 | 11261 | 11261 | 11261 | 11261 | 11261 | 8522 | 0
2017-07-21 15:30 | 11255 | 11255 | 11255 | 11255 | 11255 | 8522 | 0
2017-07-21 15:40 | 11247 | 11247 | 11247 | 11247 | 11247 | 8522 | 0
2017-07-21 15:50 | 11245 | 11245 | 11245 | 11245 | 11245 | 8522 | 0
2017-07-21 16:00 | 11246 | 11246 | 11246 | 11246 | 11246 | 8522 | 0
2017-07-21 16:10 | 11248 | 11248 | 11248 | 11248 | 11248 | 8522 | 0
2017-07-21 16:20 | 11253 | 11253 | 11253 | 11253 | 11253 | 8522 | 0
2017-07-21 16:30 | 11259 | 11259 | 11259 | 11259 | 11259 | 8522 | 0
2017-07-21 16:40 | 11266 | 11266 | 11266 | 11266 | 11266 | 8522 | 0
2017-07-21 16:50 | 11273 | 11273 | 11273 | 11273 | 11273 | 8522 | 0
2017-07-21 17:00 | 11280 | 11280 | 11280 | 11280 | 11280 | 8522 | 0
2017-07-21 17:10 | 11284 | 11284 | 11284 | 11284 | 11284 | 8522 | 0
2017-07-21 17:20 | 11287 | 11287 | 11287 | 11287 | 11287 | 8522 | 0
2017-07-21 17:30 | 11289 | 11289 | 11289 | 11289 | 11289 | 8522 | 0
2017-07-21 17:40 | 11289 | 11289 | 11289 | 11289 | 11289 | 8522 | 0
2017-07-21 17:50 | 11292 | 11292 | 11292 | 11292 | 11292 | 8522 | 0
2017-07-21 18:00 | 11296 | 11296 | 11296 | 11296 | 11296 | 8522 | 0
2017-07-21 18:10 | 11301 | 11301 | 11301 | 11301 | 11301 | 8522 | 0
2017-07-21 18:20 | 11307 | 11307 | 11307 | 11307 | 11307 | 8522 | 0
2017-07-21 18:30 | 11311 | 11311 | 11311 | 11311 | 11311 | 8522 | 0
2017-07-21 18:40 | 11314 | 11314 | 11314 | 11314 | 11314 | 8522 | 0
2017-07-21 18:50 | 11318 | 11318 | 11318 | 11318 | 11318 | 8522 | 0
2017-07-21 19:00 | 11322 | 11322 | 11322 | 11322 | 11322 | 8522 | 0
2017-07-21 19:10 | 11327 | 11327 | 11327 | 11327 | 11327 | 8522 | 0
2017-07-21 19:20 | 11332 | 11332 | 11332 | 11332 | 11332 | 8522 | 0
2017-07-21 19:30 | 11337 | 11337 | 11337 | 11337 | 11337 | 8522 | 0
2017-07-21 19:40 | 11341 | 11341 | 11341 | 11341 | 11341 | 8522 | 0
2017-07-21 19:50 | 11343 | 11343 | 11343 | 11343 | 11343 | 8522 | 0
2017-07-21 20:00 | 11342 | 11342 | 11342 | 11342 | 11342 | 8522 | 0
2017-07-21 20:10 | 11332 | 11332 | 11332 | 11332 | 11332 | 8522 | 0
2017-07-21 20:20 | 11320 | 11320 | 11320 | 11320 | 11320 | 8522 | 0
2017-07-21 20:30 | 11310 | 11310 | 11310 | 11310 | 11310 | 8522 | 0
2017-07-21 20:40 | 11269 | 11269 | 11269 | 11269 | 11269 | 8522 | 0
2017-07-21 20:50 | 11241 | 11241 | 11241 | 11241 | 11241 | 8522 | 0
2017-07-21 21:00 | 11190 | 11190 | 11190 | 11190 | 11190 | 8522 | 0
2017-07-21 21:10 | 11136 | 11136 | 11136 | 11136 | 11136 | 8522 | 0
2017-07-21 21:20 | 11086 | 11086 | 11086 | 11086 | 11086 | 8522 | 0
2017-07-21 21:30 | 11038 | 11038 | 11038 | 11038 | 11038 | 8522 | 0
2017-07-21 21:40 | 10990 | 10990 | 10990 | 10990 | 10990 | 8522 | 0
2017-07-21 21:50 | 10944 | 10944 | 10944 | 10944 | 10944 | 8522 | 0
2017-07-21 22:00 | 10900 | 10900 | 10900 | 10900 | 10900 | 8522 | 0
2017-07-21 22:10 | 10858 | 10858 | 10858 | 10858 | 10858 | 8522 | 0
2017-07-21 22:20 | 10820 | 10820 | 10820 | 10820 | 10820 | 8522 | 0
2017-07-21 22:30 | 10785 | 10785 | 10785 | 10785 | 10785 | 8522 | 0
2017-07-21 22:40 | 10754 | 10754 | 10754 | 10754 | 10754 | 8522 | 0
2017-07-21 22:50 | 10729 | 10729 | 10729 | 10729 | 10729 | 8522 | 0
2017-07-21 23:00 | 10707 | 10707 | 10707 | 10707 | 10707 | 8522 | 0
2017-07-21 23:10 | 10689 | 10689 | 10689 | 10689 | 10689 | 8522 | 0
2017-07-21 23:20 | 10675 | 10675 | 10675 | 10675 | 10675 | 8522 | 0
2017-07-21 23:30 | 10665 | 10665 | 10665 | 10665 | 10665 | 8522 | 0
2017-07-21 23:40 | 10661 | 10661 | 10661 | 10661 | 10661 | 8522 | 0
2017-07-21 23:50 | 10660 | 10660 | 10660 | 10660 | 10660 | 8522 | 0
2017-07-22 | 10663 | 10663 | 10663 | 10663 | 10663 | 8522 | 0
2017-07-22 00:10 | 10668 | 10668 | 10668 | 10668 | 10668 | 8522 | 0
2017-07-22 00:20 | 10665 | 10665 | 10665 | 10665 | 10665 | 8522 | 0
2017-07-22 00:30 | 10657 | 10657 | 10657 | 10657 | 10657 | 8522 | 0
2017-07-22 00:40 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-22 00:50 | 10646 | 10646 | 10646 | 10646 | 10646 | 8522 | 0
2017-07-22 01:00 | 10645 | 10645 | 10645 | 10645 | 10645 | 8522 | 0
2017-07-22 01:10 | 10647 | 10647 | 10647 | 10647 | 10647 | 8522 | 0
2017-07-22 01:20 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-22 01:30 | 10656 | 10656 | 10656 | 10656 | 10656 | 8522 | 0
2017-07-22 01:40 | 10665 | 10665 | 10665 | 10665 | 10665 | 8522 | 0
2017-07-22 01:50 | 10675 | 10675 | 10675 | 10675 | 10675 | 8522 | 0
2017-07-22 02:00 | 10688 | 10688 | 10688 | 10688 | 10688 | 8522 | 0
2017-07-22 02:10 | 10680 | 10680 | 10680 | 10680 | 10680 | 8522 | 0
2017-07-22 02:20 | 10662 | 10662 | 10662 | 10662 | 10662 | 8522 | 0
2017-07-22 02:30 | 10663 | 10663 | 10663 | 10663 | 10663 | 8522 | 0
2017-07-22 02:40 | 10662 | 10662 | 10662 | 10662 | 10662 | 8522 | 0
2017-07-22 02:50 | 10664 | 10664 | 10664 | 10664 | 10664 | 8522 | 0
2017-07-22 03:00 | 10666 | 10666 | 10666 | 10666 | 10666 | 8522 | 0
2017-07-22 03:10 | 10664 | 10664 | 10664 | 10664 | 10664 | 8522 | 0
2017-07-22 03:20 | 10652 | 10652 | 10652 | 10652 | 10652 | 8522 | 0
2017-07-22 03:30 | 10631 | 10631 | 10631 | 10631 | 10631 | 8522 | 0
2017-07-22 03:40 | 10601 | 10601 | 10601 | 10601 | 10601 | 8522 | 0
2017-07-22 03:50 | 10569 | 10569 | 10569 | 10569 | 10569 | 8522 | 0
2017-07-22 04:00 | 10537 | 10537 | 10537 | 10537 | 10537 | 8522 | 0
2017-07-22 04:10 | 10505 | 10505 | 10505 | 10505 | 10505 | 8522 | 0
2017-07-22 04:20 | 10474 | 10474 | 10474 | 10474 | 10474 | 8522 | 0
2017-07-22 04:30 | 10447 | 10447 | 10447 | 10447 | 10447 | 8522 | 0
2017-07-22 04:40 | 10426 | 10426 | 10426 | 10426 | 10426 | 8522 | 0
2017-07-22 04:50 | 10403 | 10403 | 10403 | 10403 | 10403 | 8522 | 0
2017-07-22 05:00 | 10376 | 10376 | 10376 | 10376 | 10376 | 8522 | 0
2017-07-22 05:10 | 10347 | 10347 | 10347 | 10347 | 10347 | 8522 | 0
2017-07-22 05:20 | 10338 | 10338 | 10338 | 10338 | 10338 | 8522 | 0
2017-07-22 05:30 | 10336 | 10336 | 10336 | 10336 | 10336 | 8522 | 0
2017-07-22 05:40 | 10340 | 10340 | 10340 | 10340 | 10340 | 8522 | 0
2017-07-22 05:50 | 10348 | 10348 | 10348 | 10348 | 10348 | 8522 | 0
2017-07-22 06:00 | 10359 | 10359 | 10359 | 10359 | 10359 | 8522 | 0
2017-07-22 06:10 | 10374 | 10374 | 10374 | 10374 | 10374 | 8522 | 0
2017-07-22 06:20 | 10392 | 10392 | 10392 | 10392 | 10392 | 8522 | 0
2017-07-22 06:30 | 10412 | 10412 | 10412 | 10412 | 10412 | 8522 | 0
2017-07-22 06:40 | 10435 | 10435 | 10435 | 10435 | 10435 | 8522 | 0
2017-07-22 06:50 | 10459 | 10459 | 10459 | 10459 | 10459 | 8522 | 0
2017-07-22 07:00 | 10485 | 10485 | 10485 | 10485 | 10485 | 8522 | 0
2017-07-22 07:10 | 10512 | 10512 | 10512 | 10512 | 10512 | 8522 | 0
2017-07-22 07:20 | 10539 | 10539 | 10539 | 10539 | 10539 | 8522 | 0
2017-07-22 07:30 | 10567 | 10567 | 10567 | 10567 | 10567 | 8522 | 0
2017-07-22 07:40 | 10596 | 10596 | 10596 | 10596 | 10596 | 8522 | 0
2017-07-22 07:50 | 10626 | 10626 | 10626 | 10626 | 10626 | 8522 | 0
2017-07-22 08:00 | 10654 | 10654 | 10654 | 10654 | 10654 | 8522 | 0
2017-07-22 08:10 | 10684 | 10684 | 10684 | 10684 | 10684 | 8522 | 0
2017-07-22 08:20 | 10714 | 10714 | 10714 | 10714 | 10714 | 8522 | 0
2017-07-22 08:30 | 10742 | 10742 | 10742 | 10742 | 10742 | 8522 | 0
2017-07-22 08:40 | 10769 | 10769 | 10769 | 10769 | 10769 | 8522 | 0
2017-07-22 08:50 | 10795 | 10795 | 10795 | 10795 | 10795 | 8522 | 0
2017-07-22 09:00 | 10819 | 10819 | 10819 | 10819 | 10819 | 8522 | 0
2017-07-22 09:10 | 10843 | 10843 | 10843 | 10843 | 10843 | 8522 | 0
2017-07-22 09:20 | 10868 | 10868 | 10868 | 10868 | 10868 | 8522 | 0
2017-07-22 09:30 | 10892 | 10892 | 10892 | 10892 | 10892 | 8522 | 0
2017-07-22 09:40 | 10917 | 10917 | 10917 | 10917 | 10917 | 8522 | 0
2017-07-22 09:50 | 10941 | 10941 | 10941 | 10941 | 10941 | 8522 | 0
2017-07-22 10:00 | 10966 | 10966 | 10966 | 10966 | 10966 | 8522 | 0
2017-07-22 10:10 | 10993 | 10993 | 10993 | 10993 | 10993 | 8522 | 0
2017-07-22 10:20 | 11020 | 11020 | 11020 | 11020 | 11020 | 8522 | 0
2017-07-22 10:30 | 11047 | 11047 | 11047 | 11047 | 11047 | 8522 | 0
2017-07-22 10:40 | 11073 | 11073 | 11073 | 11073 | 11073 | 8522 | 0
2017-07-22 10:50 | 11098 | 11098 | 11098 | 11098 | 11098 | 8522 | 0
2017-07-22 11:00 | 11123 | 11123 | 11123 | 11123 | 11123 | 8522 | 0
2017-07-22 11:10 | 11148 | 11148 | 11148 | 11148 | 11148 | 8522 | 0
2017-07-22 11:20 | 11172 | 11172 | 11172 | 11172 | 11172 | 8522 | 0
2017-07-22 11:30 | 11196 | 11196 | 11196 | 11196 | 11196 | 8522 | 0
2017-07-22 11:40 | 11220 | 11220 | 11220 | 11220 | 11220 | 8522 | 0
2017-07-22 11:50 | 11245 | 11245 | 11245 | 11245 | 11245 | 8522 | 0
2017-07-22 12:00 | 11269 | 11269 | 11269 | 11269 | 11269 | 8522 | 0
2017-07-22 12:10 | 11292 | 11292 | 11292 | 11292 | 11292 | 8522 | 0
2017-07-22 12:20 | 11315 | 11315 | 11315 | 11315 | 11315 | 8522 | 0
2017-07-22 12:30 | 11338 | 11338 | 11338 | 11338 | 11338 | 8522 | 0
2017-07-22 12:40 | 11361 | 11361 | 11361 | 11361 | 11361 | 8522 | 0
2017-07-22 12:50 | 11386 | 11386 | 11386 | 11386 | 11386 | 8522 | 0
2017-07-22 13:00 | 11412 | 11412 | 11412 | 11412 | 11412 | 8522 | 0
2017-07-22 13:10 | 11440 | 11440 | 11440 | 11440 | 11440 | 8522 | 0
2017-07-22 13:20 | 11469 | 11469 | 11469 | 11469 | 11469 | 8522 | 0
2017-07-22 13:30 | 11499 | 11499 | 11499 | 11499 | 11499 | 8522 | 0
2017-07-22 13:40 | 11529 | 11529 | 11529 | 11529 | 11529 | 8522 | 0
2017-07-22 13:50 | 11559 | 11559 | 11559 | 11559 | 11559 | 8522 | 0
2017-07-22 14:00 | 11589 | 11589 | 11589 | 11589 | 11589 | 8522 | 0
2017-07-22 14:10 | 11618 | 11618 | 11618 | 11618 | 11618 | 8522 | 0
2017-07-22 14:20 | 11647 | 11647 | 11647 | 11647 | 11647 | 8522 | 0
2017-07-22 14:30 | 11664 | 11664 | 11664 | 11664 | 11664 | 8522 | 0
2017-07-22 14:40 | 11668 | 11668 | 11668 | 11668 | 11668 | 8522 | 0
2017-07-22 14:50 | 11679 | 11679 | 11679 | 11679 | 11679 | 8522 | 0
2017-07-22 15:00 | 11692 | 11692 | 11692 | 11692 | 11692 | 8522 | 0
2017-07-22 15:10 | 11705 | 11705 | 11705 | 11705 | 11705 | 8522 | 0
2017-07-22 15:20 | 11716 | 11716 | 11716 | 11716 | 11716 | 8522 | 0
2017-07-22 15:30 | 11726 | 11726 | 11726 | 11726 | 11726 | 8522 | 0
2017-07-22 15:40 | 11732 | 11732 | 11732 | 11732 | 11732 | 8522 | 0
2017-07-22 15:50 | 11723 | 11723 | 11723 | 11723 | 11723 | 8522 | 0
2017-07-22 16:00 | 11709 | 11709 | 11709 | 11709 | 11709 | 8522 | 0
2017-07-22 16:10 | 11702 | 11702 | 11702 | 11702 | 11702 | 8522 | 0
2017-07-22 16:20 | 11701 | 11701 | 11701 | 11701 | 11701 | 8522 | 0
2017-07-22 16:30 | 11699 | 11699 | 11699 | 11699 | 11699 | 8522 | 0
2017-07-22 16:40 | 11655 | 11655 | 11655 | 11655 | 11655 | 8522 | 0
2017-07-22 16:50 | 11631 | 11631 | 11631 | 11631 | 11631 | 8522 | 0
2017-07-22 17:00 | 11607 | 11607 | 11607 | 11607 | 11607 | 8522 | 0
2017-07-22 17:10 | 11580 | 11580 | 11580 | 11580 | 11580 | 8522 | 0
2017-07-22 17:20 | 11543 | 11543 | 11543 | 11543 | 11543 | 8522 | 0
2017-07-22 17:30 | 11496 | 11496 | 11496 | 11496 | 11496 | 8522 | 0
2017-07-22 17:40 | 11447 | 11447 | 11447 | 11447 | 11447 | 8522 | 0
2017-07-22 17:50 | 11399 | 11399 | 11399 | 11399 | 11399 | 8522 | 0
2017-07-22 18:00 | 11341 | 11341 | 11341 | 11341 | 11341 | 8522 | 0
2017-07-22 18:10 | 11291 | 11291 | 11291 | 11291 | 11291 | 8522 | 0
2017-07-22 18:20 | 11248 | 11248 | 11248 | 11248 | 11248 | 8522 | 0
2017-07-22 18:30 | 11207 | 11207 | 11207 | 11207 | 11207 | 8522 | 0
2017-07-22 18:40 | 11167 | 11167 | 11167 | 11167 | 11167 | 8522 | 0
2017-07-22 18:50 | 11123 | 11123 | 11123 | 11123 | 11123 | 8522 | 0
2017-07-22 19:00 | 11066 | 11066 | 11066 | 11066 | 11066 | 8522 | 0
2017-07-22 19:10 | 11018 | 11018 | 11018 | 11018 | 11018 | 8522 | 0
2017-07-22 19:20 | 10970 | 10970 | 10970 | 10970 | 10970 | 8522 | 0
2017-07-22 19:30 | 10925 | 10925 | 10925 | 10925 | 10925 | 8522 | 0
2017-07-22 19:40 | 10887 | 10887 | 10887 | 10887 | 10887 | 8522 | 0
2017-07-22 19:50 | 10851 | 10851 | 10851 | 10851 | 10851 | 8522 | 0
2017-07-22 20:00 | 10815 | 10815 | 10815 | 10815 | 10815 | 8522 | 0
2017-07-22 20:10 | 10779 | 10779 | 10779 | 10779 | 10779 | 8522 | 0
2017-07-22 20:20 | 10740 | 10740 | 10740 | 10740 | 10740 | 8522 | 0
2017-07-22 20:30 | 10705 | 10705 | 10705 | 10705 | 10705 | 8522 | 0
2017-07-22 20:40 | 10675 | 10675 | 10675 | 10675 | 10675 | 8522 | 0
2017-07-22 20:50 | 10649 | 10649 | 10649 | 10649 | 10649 | 8522 | 0
2017-07-22 21:00 | 10628 | 10628 | 10628 | 10628 | 10628 | 8522 | 0
2017-07-22 21:10 | 10609 | 10609 | 10609 | 10609 | 10609 | 8522 | 0
2017-07-22 21:20 | 10592 | 10592 | 10592 | 10592 | 10592 | 8522 | 0
2017-07-22 21:30 | 10571 | 10571 | 10571 | 10571 | 10571 | 8522 | 0
2017-07-22 21:40 | 10551 | 10551 | 10551 | 10551 | 10551 | 8522 | 0
2017-07-22 21:50 | 10528 | 10528 | 10528 | 10528 | 10528 | 8522 | 0
2017-07-22 22:00 | 10508 | 10508 | 10508 | 10508 | 10508 | 8522 | 0
2017-07-22 22:10 | 10490 | 10490 | 10490 | 10490 | 10490 | 8522 | 0
2017-07-22 22:20 | 10474 | 10474 | 10474 | 10474 | 10474 | 8522 | 0
2017-07-22 22:30 | 10461 | 10461 | 10461 | 10461 | 10461 | 8522 | 0
2017-07-22 22:40 | 10451 | 10451 | 10451 | 10451 | 10451 | 8522 | 0
2017-07-22 22:50 | 10439 | 10439 | 10439 | 10439 | 10439 | 8522 | 0
2017-07-22 23:00 | 10428 | 10428 | 10428 | 10428 | 10428 | 8522 | 0
2017-07-22 23:10 | 10419 | 10419 | 10419 | 10419 | 10419 | 8522 | 0
2017-07-22 23:20 | 10411 | 10411 | 10411 | 10411 | 10411 | 8522 | 0
2017-07-22 23:30 | 10405 | 10405 | 10405 | 10405 | 10405 | 8522 | 0
2017-07-22 23:40 | 10399 | 10399 | 10399 | 10399 | 10399 | 8522 | 0
2017-07-22 23:50 | 10395 | 10395 | 10395 | 10395 | 10395 | 8522 | 0
2017-07-23 | 10388 | 10388 | 10388 | 10388 | 10388 | 8522 | 0
2017-07-23 00:10 | 10375 | 10375 | 10375 | 10375 | 10375 | 8522 | 0
2017-07-23 00:20 | 10358 | 10358 | 10358 | 10358 | 10358 | 8522 | 0
2017-07-23 00:30 | 10342 | 10342 | 10342 | 10342 | 10342 | 8522 | 0
2017-07-23 00:40 | 10330 | 10330 | 10330 | 10330 | 10330 | 8522 | 0
2017-07-23 00:50 | 10324 | 10324 | 10324 | 10324 | 10324 | 8522 | 0
2017-07-23 01:00 | 10322 | 10322 | 10322 | 10322 | 10322 | 8522 | 0
2017-07-23 01:10 | 10324 | 10324 | 10324 | 10324 | 10324 | 8522 | 0
2017-07-23 01:20 | 10332 | 10332 | 10332 | 10332 | 10332 | 8522 | 0
2017-07-23 01:30 | 10344 | 10344 | 10344 | 10344 | 10344 | 8522 | 0
2017-07-23 01:40 | 10360 | 10360 | 10360 | 10360 | 10360 | 8522 | 0
2017-07-23 01:50 | 10371 | 10371 | 10371 | 10371 | 10371 | 8522 | 0
2017-07-23 02:00 | 10356 | 10356 | 10356 | 10356 | 10356 | 8522 | 0
2017-07-23 02:10 | 10346 | 10346 | 10346 | 10346 | 10346 | 8522 | 0
2017-07-23 02:20 | 10348 | 10348 | 10348 | 10348 | 10348 | 8522 | 0
2017-07-23 02:30 | 10352 | 10352 | 10352 | 10352 | 10352 | 8522 | 0
2017-07-23 02:40 | 10357 | 10357 | 10357 | 10357 | 10357 | 8522 | 0
2017-07-23 02:50 | 10363 | 10363 | 10363 | 10363 | 10363 | 8522 | 0
2017-07-23 03:00 | 10367 | 10367 | 10367 | 10367 | 10367 | 8522 | 0
2017-07-23 03:10 | 10372 | 10372 | 10372 | 10372 | 10372 | 8522 | 0
2017-07-23 03:20 | 10385 | 10385 | 10385 | 10385 | 10385 | 8522 | 0
2017-07-23 03:30 | 10394 | 10394 | 10394 | 10394 | 10394 | 8522 | 0
2017-07-23 03:40 | 10397 | 10397 | 10397 | 10397 | 10397 | 8522 | 0
2017-07-23 03:50 | 10401 | 10401 | 10401 | 10401 | 10401 | 8522 | 0
2017-07-23 04:00 | 10414 | 10414 | 10414 | 10414 | 10414 | 8522 | 0
2017-07-23 04:10 | 10434 | 10434 | 10434 | 10434 | 10434 | 8522 | 0
2017-07-23 04:20 | 10456 | 10456 | 10456 | 10456 | 10456 | 8522 | 0
2017-07-23 04:30 | 10478 | 10478 | 10478 | 10478 | 10478 | 8522 | 0
2017-07-23 04:40 | 10498 | 10498 | 10498 | 10498 | 10498 | 8522 | 0
2017-07-23 04:50 | 10517 | 10517 | 10517 | 10517 | 10517 | 8522 | 0
2017-07-23 05:00 | 10534 | 10534 | 10534 | 10534 | 10534 | 8522 | 0
2017-07-23 05:10 | 10549 | 10549 | 10549 | 10549 | 10549 | 8522 | 0
2017-07-23 05:20 | 10560 | 10560 | 10560 | 10560 | 10560 | 8522 | 0
2017-07-23 05:30 | 10569 | 10569 | 10569 | 10569 | 10569 | 8522 | 0
2017-07-23 05:40 | 10560 | 10560 | 10560 | 10560 | 10560 | 8522 | 0
2017-07-23 05:50 | 10535 | 10535 | 10535 | 10535 | 10535 | 8522 | 0
2017-07-23 06:00 | 10513 | 10513 | 10513 | 10513 | 10513 | 8522 | 0
2017-07-23 06:10 | 10494 | 10494 | 10494 | 10494 | 10494 | 8522 | 0
2017-07-23 06:20 | 10486 | 10486 | 10486 | 10486 | 10486 | 8522 | 0
2017-07-23 06:30 | 10489 | 10489 | 10489 | 10489 | 10489 | 8522 | 0
2017-07-23 06:40 | 10500 | 10500 | 10500 | 10500 | 10500 | 8522 | 0
2017-07-23 06:50 | 10518 | 10518 | 10518 | 10518 | 10518 | 8522 | 0
2017-07-23 07:00 | 10540 | 10540 | 10540 | 10540 | 10540 | 8522 | 0
2017-07-23 07:10 | 10565 | 10565 | 10565 | 10565 | 10565 | 8522 | 0
2017-07-23 07:20 | 10592 | 10592 | 10592 | 10592 | 10592 | 8522 | 0
2017-07-23 07:30 | 10620 | 10620 | 10620 | 10620 | 10620 | 8522 | 0
2017-07-23 07:40 | 10649 | 10649 | 10649 | 10649 | 10649 | 8522 | 0
2017-07-23 07:50 | 10678 | 10678 | 10678 | 10678 | 10678 | 8522 | 0
2017-07-23 08:00 | 10707 | 10707 | 10707 | 10707 | 10707 | 8522 | 0
2017-07-23 08:10 | 10735 | 10735 | 10735 | 10735 | 10735 | 8522 | 0
2017-07-23 08:20 | 10763 | 10763 | 10763 | 10763 | 10763 | 8522 | 0
2017-07-23 08:30 | 10790 | 10790 | 10790 | 10790 | 10790 | 8522 | 0
2017-07-23 08:40 | 10817 | 10817 | 10817 | 10817 | 10817 | 8522 | 0
2017-07-23 08:50 | 10843 | 10843 | 10843 | 10843 | 10843 | 8522 | 0
2017-07-23 09:00 | 10869 | 10869 | 10869 | 10869 | 10869 | 8522 | 0
2017-07-23 09:10 | 10894 | 10894 | 10894 | 10894 | 10894 | 8522 | 0
2017-07-23 09:20 | 10919 | 10919 | 10919 | 10919 | 10919 | 8522 | 0
2017-07-23 09:30 | 10943 | 10943 | 10943 | 10943 | 10943 | 8522 | 0
2017-07-23 09:40 | 10966 | 10966 | 10966 | 10966 | 10966 | 8522 | 0
2017-07-23 09:50 | 10989 | 10989 | 10989 | 10989 | 10989 | 8522 | 0
2017-07-23 10:00 | 11011 | 11011 | 11011 | 11011 | 11011 | 8522 | 0
2017-07-23 10:10 | 11034 | 11034 | 11034 | 11034 | 11034 | 8522 | 0
2017-07-23 10:20 | 11057 | 11057 | 11057 | 11057 | 11057 | 8522 | 0
2017-07-23 10:30 | 11080 | 11080 | 11080 | 11080 | 11080 | 8522 | 0
2017-07-23 10:40 | 11102 | 11102 | 11102 | 11102 | 11102 | 8522 | 0
2017-07-23 10:50 | 11124 | 11124 | 11124 | 11124 | 11124 | 8522 | 0
2017-07-23 11:00 | 11145 | 11145 | 11145 | 11145 | 11145 | 8522 | 0
2017-07-23 11:10 | 11165 | 11165 | 11165 | 11165 | 11165 | 8522 | 0
2017-07-23 11:20 | 11185 | 11185 | 11185 | 11185 | 11185 | 8522 | 0
2017-07-23 11:30 | 11203 | 11203 | 11203 | 11203 | 11203 | 8522 | 0
2017-07-23 11:40 | 11220 | 11220 | 11220 | 11220 | 11220 | 8522 | 0
2017-07-23 11:50 | 11236 | 11236 | 11236 | 11236 | 11236 | 8522 | 0
2017-07-23 12:00 | 11250 | 11250 | 11250 | 11250 | 11250 | 8522 | 0
2017-07-23 12:10 | 11262 | 11262 | 11262 | 11262 | 11262 | 8522 | 0
2017-07-23 12:20 | 11273 | 11273 | 11273 | 11273 | 11273 | 8522 | 0
2017-07-23 12:30 | 11284 | 11284 | 11284 | 11284 | 11284 | 8522 | 0
2017-07-23 12:40 | 11294 | 11294 | 11294 | 11294 | 11294 | 8522 | 0
2017-07-23 12:50 | 11303 | 11303 | 11303 | 11303 | 11303 | 8522 | 0
2017-07-23 13:00 | 11312 | 11312 | 11312 | 11312 | 11312 | 8522 | 0
2017-07-23 13:10 | 11322 | 11322 | 11322 | 11322 | 11322 | 8522 | 0
2017-07-23 13:20 | 11333 | 11333 | 11333 | 11333 | 11333 | 8522 | 0
2017-07-23 13:30 | 11345 | 11345 | 11345 | 11345 | 11345 | 8522 | 0
2017-07-23 13:40 | 11357 | 11357 | 11357 | 11357 | 11357 | 8522 | 0
2017-07-23 13:50 | 11369 | 11369 | 11369 | 11369 | 11369 | 8522 | 0
2017-07-23 14:00 | 11381 | 11381 | 11381 | 11381 | 11381 | 8522 | 0
2017-07-23 14:10 | 11394 | 11394 | 11394 | 11394 | 11394 | 8522 | 0
2017-07-23 14:20 | 11408 | 11408 | 11408 | 11408 | 11408 | 8522 | 0
2017-07-23 14:30 | 11423 | 11423 | 11423 | 11423 | 11423 | 8522 | 0
2017-07-23 14:40 | 11438 | 11438 | 11438 | 11438 | 11438 | 8522 | 0
2017-07-23 14:50 | 11453 | 11453 | 11453 | 11453 | 11453 | 8522 | 0
2017-07-23 15:00 | 11465 | 11465 | 11465 | 11465 | 11465 | 8522 | 0
2017-07-23 15:10 | 11472 | 11472 | 11472 | 11472 | 11472 | 8522 | 0
2017-07-23 15:20 | 11476 | 11476 | 11476 | 11476 | 11476 | 8522 | 0
2017-07-23 15:30 | 11479 | 11479 | 11479 | 11479 | 11479 | 8522 | 0
2017-07-23 15:40 | 11483 | 11483 | 11483 | 11483 | 11483 | 8522 | 0
2017-07-23 15:50 | 11488 | 11488 | 11488 | 11488 | 11488 | 8522 | 0
2017-07-23 16:00 | 11493 | 11493 | 11493 | 11493 | 11493 | 8522 | 0
2017-07-23 16:10 | 11498 | 11498 | 11498 | 11498 | 11498 | 8522 | 0
2017-07-23 16:20 | 11500 | 11500 | 11500 | 11500 | 11500 | 8522 | 0
2017-07-23 16:30 | 11497 | 11497 | 11497 | 11497 | 11497 | 8522 | 0
2017-07-23 16:40 | 11489 | 11489 | 11489 | 11489 | 11489 | 8522 | 0
2017-07-23 16:50 | 11473 | 11473 | 11473 | 11473 | 11473 | 8522 | 0
2017-07-23 17:00 | 11454 | 11454 | 11454 | 11454 | 11454 | 8522 | 0
2017-07-23 17:10 | 11430 | 11430 | 11430 | 11430 | 11430 | 8522 | 0
2017-07-23 17:20 | 11408 | 11408 | 11408 | 11408 | 11408 | 8522 | 0
2017-07-23 17:30 | 11391 | 11391 | 11391 | 11391 | 11391 | 8522 | 0
2017-07-23 17:40 | 11374 | 11374 | 11374 | 11374 | 11374 | 8522 | 0
2017-07-23 17:50 | 11356 | 11356 | 11356 | 11356 | 11356 | 8522 | 0
2017-07-23 18:00 | 11315 | 11315 | 11315 | 11315 | 11315 | 8522 | 0
2017-07-23 18:10 | 11269 | 11269 | 11269 | 11269 | 11269 | 8522 | 0
2017-07-23 18:20 | 11223 | 11223 | 11223 | 11223 | 11223 | 8522 | 0
2017-07-23 18:30 | 11169 | 11169 | 11169 | 11169 | 11169 | 8522 | 0
2017-07-23 18:40 | 11120 | 11120 | 11120 | 11120 | 11120 | 8522 | 0
2017-07-23 18:50 | 11068 | 11068 | 11068 | 11068 | 11068 | 8522 | 0
2017-07-23 19:00 | 11009 | 11009 | 11009 | 11009 | 11009 | 8522 | 0
2017-07-23 19:10 | 10957 | 10957 | 10957 | 10957 | 10957 | 8522 | 0
2017-07-23 19:20 | 10909 | 10909 | 10909 | 10909 | 10909 | 8522 | 0
2017-07-23 19:30 | 10864 | 10864 | 10864 | 10864 | 10864 | 8522 | 0
2017-07-23 19:40 | 10822 | 10822 | 10822 | 10822 | 10822 | 8522 | 0
2017-07-23 19:50 | 10785 | 10785 | 10785 | 10785 | 10785 | 8522 | 0
2017-07-23 20:00 | 10751 | 10751 | 10751 | 10751 | 10751 | 8522 | 0
2017-07-23 20:10 | 10716 | 10716 | 10716 | 10716 | 10716 | 8522 | 0
2017-07-23 20:20 | 10678 | 10678 | 10678 | 10678 | 10678 | 8522 | 0
2017-07-23 20:30 | 10639 | 10639 | 10639 | 10639 | 10639 | 8522 | 0
2017-07-23 20:40 | 10599 | 10599 | 10599 | 10599 | 10599 | 8522 | 0
2017-07-23 20:50 | 10548 | 10548 | 10548 | 10548 | 10548 | 8522 | 0
2017-07-23 21:00 | 10481 | 10481 | 10481 | 10481 | 10481 | 8522 | 0
2017-07-23 21:10 | 10430 | 10430 | 10430 | 10430 | 10430 | 8522 | 0
2017-07-23 21:20 | 10399 | 10399 | 10399 | 10399 | 10399 | 8522 | 0
2017-07-23 21:30 | 10378 | 10378 | 10378 | 10378 | 10378 | 8522 | 0
2017-07-23 21:40 | 10381 | 10381 | 10381 | 10381 | 10381 | 8522 | 0
2017-07-23 21:50 | 10394 | 10394 | 10394 | 10394 | 10394 | 8522 | 0
2017-07-23 22:00 | 10408 | 10408 | 10408 | 10408 | 10408 | 8522 | 0
2017-07-23 22:10 | 10418 | 10418 | 10418 | 10418 | 10418 | 8522 | 0
2017-07-23 22:20 | 10402 | 10402 | 10402 | 10402 | 10402 | 8522 | 0
2017-07-23 22:30 | 10395 | 10395 | 10395 | 10395 | 10395 | 8522 | 0
2017-07-23 22:40 | 10409 | 10409 | 10409 | 10409 | 10409 | 8522 | 0
2017-07-23 22:50 | 10434 | 10434 | 10434 | 10434 | 10434 | 8522 | 0
2017-07-23 23:00 | 10463 | 10463 | 10463 | 10463 | 10463 | 8522 | 0
2017-07-23 23:10 | 10491 | 10491 | 10491 | 10491 | 10491 | 8522 | 0
2017-07-23 23:20 | 10517 | 10517 | 10517 | 10517 | 10517 | 8522 | 0
2017-07-23 23:30 | 10541 | 10541 | 10541 | 10541 | 10541 | 8522 | 0
2017-07-23 23:40 | 10563 | 10563 | 10563 | 10563 | 10563 | 8522 | 0
2017-07-23 23:50 | 10569 | 10569 | 10569 | 10569 | 10569 | 8522 | 0
2017-07-24 | 10560 | 10560 | 10560 | 10560 | 10560 | 8522 | 0
2017-07-24 00:10 | 10537 | 10537 | 10537 | 10537 | 10537 | 8522 | 0
2017-07-24 00:20 | 10529 | 10529 | 10529 | 10529 | 10529 | 8522 | 0
2017-07-24 00:30 | 10538 | 10538 | 10538 | 10538 | 10538 | 8522 | 0
2017-07-24 00:40 | 10555 | 10555 | 10555 | 10555 | 10555 | 8522 | 0
2017-07-24 00:50 | 10574 | 10574 | 10574 | 10574 | 10574 | 8522 | 0
2017-07-24 01:00 | 10592 | 10592 | 10592 | 10592 | 10592 | 8522 | 0
2017-07-24 01:10 | 10586 | 10586 | 10586 | 10586 | 10586 | 8522 | 0
2017-07-24 01:20 | 10573 | 10573 | 10573 | 10573 | 10573 | 8522 | 0
2017-07-24 01:30 | 10563 | 10563 | 10563 | 10563 | 10563 | 8522 | 0
2017-07-24 01:40 | 10559 | 10559 | 10559 | 10559 | 10559 | 8522 | 0
2017-07-24 01:50 | 10567 | 10567 | 10567 | 10567 | 10567 | 8522 | 0
2017-07-24 02:00 | 10598 | 10598 | 10598 | 10598 | 10598 | 8522 | 0
2017-07-24 02:10 | 10640 | 10640 | 10640 | 10640 | 10640 | 8522 | 0
2017-07-24 02:20 | 10684 | 10684 | 10684 | 10684 | 10684 | 8522 | 0
2017-07-24 02:30 | 10728 | 10728 | 10728 | 10728 | 10728 | 8522 | 0
2017-07-24 02:40 | 10769 | 10769 | 10769 | 10769 | 10769 | 8522 | 0
2017-07-24 02:50 | 10791 | 10791 | 10791 | 10791 | 10791 | 8522 | 0
2017-07-24 03:00 | 10806 | 10806 | 10806 | 10806 | 10806 | 8522 | 0
2017-07-24 03:10 | 10811 | 10811 | 10811 | 10811 | 10811 | 8522 | 0
2017-07-24 03:20 | 10815 | 10815 | 10815 | 10815 | 10815 | 8522 | 0
2017-07-24 03:30 | 10826 | 10826 | 10826 | 10826 | 10826 | 8522 | 0
2017-07-24 03:40 | 10856 | 10856 | 10856 | 10856 | 10856 | 8522 | 0
2017-07-24 03:50 | 10897 | 10897 | 10897 | 10897 | 10897 | 8522 | 0
2017-07-24 04:00 | 10933 | 10933 | 10933 | 10933 | 10933 | 8522 | 0
2017-07-24 04:10 | 10955 | 10955 | 10955 | 10955 | 10955 | 8522 | 0
2017-07-24 04:20 | 10983 | 10983 | 10983 | 10983 | 10983 | 8522 | 0
2017-07-24 04:30 | 11020 | 11020 | 11020 | 11020 | 11020 | 8522 | 0
2017-07-24 04:40 | 11066 | 11066 | 11066 | 11066 | 11066 | 8522 | 0
2017-07-24 04:50 | 11117 | 11117 | 11117 | 11117 | 11117 | 8522 | 0
2017-07-24 05:00 | 11169 | 11169 | 11169 | 11169 | 11169 | 8522 | 0
2017-07-24 05:10 | 11217 | 11217 | 11217 | 11217 | 11217 | 8522 | 0
2017-07-24 05:20 | 11256 | 11256 | 11256 | 11256 | 11256 | 8522 | 0
2017-07-24 05:30 | 11289 | 11289 | 11289 | 11289 | 11289 | 8522 | 0
2017-07-24 05:40 | 11323 | 11323 | 11323 | 11323 | 11323 | 8522 | 0
2017-07-24 05:50 | 11361 | 11361 | 11361 | 11361 | 11361 | 8522 | 0
2017-07-24 06:00 | 11401 | 11401 | 11401 | 11401 | 11401 | 8522 | 0
2017-07-24 06:10 | 11442 | 11442 | 11442 | 11442 | 11442 | 8522 | 0
2017-07-24 06:20 | 11481 | 11481 | 11481 | 11481 | 11481 | 8522 | 0
2017-07-24 06:30 | 11521 | 11521 | 11521 | 11521 | 11521 | 8522 | 0
2017-07-24 06:40 | 11559 | 11559 | 11559 | 11559 | 11559 | 8522 | 0
2017-07-24 06:50 | 11598 | 11598 | 11598 | 11598 | 11598 | 8522 | 0
2017-07-24 07:00 | 11637 | 11637 | 11637 | 11637 | 11637 | 8522 | 0
2017-07-24 07:10 | 11677 | 11677 | 11677 | 11677 | 11677 | 8522 | 0
2017-07-24 07:20 | 11718 | 11718 | 11718 | 11718 | 11718 | 8522 | 0
2017-07-24 07:30 | 11760 | 11760 | 11760 | 11760 | 11760 | 8522 | 0
2017-07-24 07:40 | 11802 | 11802 | 11802 | 11802 | 11802 | 8522 | 0
2017-07-24 07:50 | 11844 | 11844 | 11844 | 11844 | 11844 | 8522 | 0
2017-07-24 08:00 | 11885 | 11885 | 11885 | 11885 | 11885 | 8522 | 0
2017-07-24 08:10 | 11926 | 11926 | 11926 | 11926 | 11926 | 8522 | 0
2017-07-24 08:20 | 11966 | 11966 | 11966 | 11966 | 11966 | 8522 | 0
2017-07-24 08:30 | 12005 | 12005 | 12005 | 12005 | 12005 | 8522 | 0
2017-07-24 08:40 | 12043 | 12043 | 12043 | 12043 | 12043 | 8522 | 0
2017-07-24 08:50 | 12080 | 12080 | 12080 | 12080 | 12080 | 8522 | 0
2017-07-24 09:00 | 12116 | 12116 | 12116 | 12116 | 12116 | 8522 | 0
2017-07-24 09:10 | 12151 | 12151 | 12151 | 12151 | 12151 | 8522 | 0
2017-07-24 09:20 | 12187 | 12187 | 12187 | 12187 | 12187 | 8522 | 0
2017-07-24 09:30 | 12221 | 12221 | 12221 | 12221 | 12221 | 8522 | 0
2017-07-24 09:40 | 12254 | 12254 | 12254 | 12254 | 12254 | 8522 | 0
2017-07-24 09:50 | 12286 | 12286 | 12286 | 12286 | 12286 | 8522 | 0
2017-07-24 10:00 | 12317 | 12317 | 12317 | 12317 | 12317 | 8523 | 0
2017-07-24 10:10 | 12348 | 12348 | 12348 | 12348 | 12348 | 8522 | 0
2017-07-24 10:20 | 12377 | 12377 | 12377 | 12377 | 12377 | 8522 | 0
2017-07-24 10:30 | 12405 | 12405 | 12405 | 12405 | 12405 | 8522 | 0
2017-07-24 10:40 | 12433 | 12433 | 12433 | 12433 | 12433 | 8522 | 0
2017-07-24 10:50 | 12462 | 12462 | 12462 | 12462 | 12462 | 8522 | 0
2017-07-24 11:00 | 12490 | 12490 | 12490 | 12490 | 12490 | 8522 | 0
2017-07-24 11:10 | 12517 | 12517 | 12517 | 12517 | 12517 | 8522 | 0
2017-07-24 11:20 | 12544 | 12544 | 12544 | 12544 | 12544 | 8522 | 0
2017-07-24 11:30 | 12570 | 12570 | 12570 | 12570 | 12570 | 8523 | 0
2017-07-24 11:40 | 12596 | 12596 | 12596 | 12596 | 12596 | 8522 | 0
2017-07-24 11:50 | 12620 | 12620 | 12620 | 12620 | 12620 | 8522 | 0
2017-07-24 12:00 | 12644 | 12644 | 12644 | 12644 | 12644 | 8522 | 0
2017-07-24 12:10 | 12667 | 12667 | 12667 | 12667 | 12667 | 8522 | 0
2017-07-24 12:20 | 12689 | 12689 | 12689 | 12689 | 12689 | 8522 | 0
2017-07-24 12:30 | 12711 | 12711 | 12711 | 12711 | 12711 | 8523 | 0
2017-07-24 12:40 | 12733 | 12733 | 12733 | 12733 | 12733 | 8522 | 0
2017-07-24 12:50 | 12754 | 12754 | 12754 | 12754 | 12754 | 8522 | 0
2017-07-24 13:00 | 12773 | 12773 | 12773 | 12773 | 12773 | 8522 | 0
2017-07-24 13:10 | 12791 | 12791 | 12791 | 12791 | 12791 | 8522 | 0
2017-07-24 13:20 | 12809 | 12809 | 12809 | 12809 | 12809 | 8522 | 0
2017-07-24 13:30 | 12827 | 12827 | 12827 | 12827 | 12827 | 8522 | 0
2017-07-24 13:40 | 12844 | 12844 | 12844 | 12844 | 12844 | 8523 | 0
2017-07-24 13:50 | 12857 | 12857 | 12857 | 12857 | 12857 | 8523 | 0
2017-07-24 14:00 | 12867 | 12867 | 12867 | 12867 | 12867 | 8523 | 0
2017-07-24 14:10 | 12875 | 12875 | 12875 | 12875 | 12875 | 8522 | 0
2017-07-24 14:20 | 12879 | 12879 | 12879 | 12879 | 12879 | 8522 | 0
2017-07-24 14:30 | 12882 | 12882 | 12882 | 12882 | 12882 | 8523 | 0
2017-07-24 14:40 | 12884 | 12884 | 12884 | 12884 | 12884 | 8522 | 0
2017-07-24 14:50 | 12884 | 12884 | 12884 | 12884 | 12884 | 8522 | 0
2017-07-24 15:00 | 12883 | 12883 | 12883 | 12883 | 12883 | 8522 | 0
2017-07-24 15:10 | 12852 | 12852 | 12852 | 12852 | 12852 | 8523 | 0
2017-07-24 15:20 | 12820 | 12820 | 12820 | 12820 | 12820 | 8523 | 0
2017-07-24 15:30 | 12800 | 12800 | 12800 | 12800 | 12800 | 8523 | 0
2017-07-24 15:40 | 12767 | 12767 | 12767 | 12767 | 12767 | 8522 | 0
2017-07-24 15:50 | 12730 | 12730 | 12730 | 12730 | 12730 | 8522 | 0
2017-07-24 16:00 | 12693 | 12693 | 12693 | 12693 | 12693 | 8522 | 0
2017-07-24 16:10 | 12654 | 12654 | 12654 | 12654 | 12654 | 8522 | 0
2017-07-24 16:20 | 12618 | 12618 | 12618 | 12618 | 12618 | 8523 | 0
2017-07-24 16:30 | 12583 | 12583 | 12583 | 12583 | 12583 | 8522 | 0
2017-07-24 16:40 | 12540 | 12540 | 12540 | 12540 | 12540 | 8522 | 0
2017-07-24 16:50 | 12509 | 12509 | 12509 | 12509 | 12509 | 8522 | 0
2017-07-24 17:00 | 12482 | 12482 | 12482 | 12482 | 12482 | 8522 | 0
2017-07-24 17:10 | 12455 | 12455 | 12455 | 12455 | 12455 | 8522 | 0
2017-07-24 17:20 | 12428 | 12428 | 12428 | 12428 | 12428 | 8522 | 0
2017-07-24 17:30 | 12397 | 12397 | 12397 | 12397 | 12397 | 8522 | 0
2017-07-24 17:40 | 12364 | 12364 | 12364 | 12364 | 12364 | 8523 | 0
2017-07-24 17:50 | 12336 | 12336 | 12336 | 12336 | 12336 | 8522 | 0
2017-07-24 18:00 | 12312 | 12312 | 12312 | 12312 | 12312 | 8522 | 0
2017-07-24 18:10 | 12292 | 12292 | 12292 | 12292 | 12292 | 8522 | 0
2017-07-24 18:20 | 12274 | 12274 | 12274 | 12274 | 12274 | 8522 | 0
2017-07-24 18:30 | 12259 | 12259 | 12259 | 12259 | 12259 | 8522 | 0
2017-07-24 18:40 | 12246 | 12246 | 12246 | 12246 | 12246 | 8522 | 0
2017-07-24 18:50 | 12193 | 12193 | 12193 | 12193 | 12193 | 8522 | 0
2017-07-24 19:00 | 12117 | 12117 | 12117 | 12117 | 12117 | 8522 | 0
2017-07-24 19:10 | 12077 | 12077 | 12077 | 12077 | 12077 | 8522 | 0
2017-07-24 19:20 | 12057 | 12057 | 12057 | 12057 | 12057 | 8522 | 0
2017-07-24 19:30 | 12045 | 12045 | 12045 | 12045 | 12045 | 8522 | 0
2017-07-24 19:40 | 12037 | 12037 | 12037 | 12037 | 12037 | 8522 | 0
2017-07-24 19:50 | 12030 | 12030 | 12030 | 12030 | 12030 | 8522 | 0
2017-07-24 20:00 | 12003 | 12003 | 12003 | 12003 | 12003 | 8522 | 0
2017-07-24 20:10 | 11966 | 11966 | 11966 | 11966 | 11966 | 8522 | 0
2017-07-24 20:20 | 11928 | 11928 | 11928 | 11928 | 11928 | 8522 | 0
2017-07-24 20:30 | 11894 | 11894 | 11894 | 11894 | 11894 | 8522 | 0
2017-07-24 20:40 | 11855 | 11855 | 11855 | 11855 | 11855 | 8522 | 0
2017-07-24 20:50 | 11823 | 11823 | 11823 | 11823 | 11823 | 8523 | 0
2017-07-24 21:00 | 11816 | 11816 | 11816 | 11816 | 11816 | 8522 | 0
2017-07-24 21:10 | 11822 | 11822 | 11822 | 11822 | 11822 | 8522 | 0
2017-07-24 21:20 | 11831 | 11831 | 11831 | 11831 | 11831 | 8522 | 0
2017-07-24 21:30 | 11838 | 11838 | 11838 | 11838 | 11838 | 8522 | 0
2017-07-24 21:40 | 11820 | 11820 | 11820 | 11820 | 11820 | 8522 | 0
2017-07-24 21:50 | 11791 | 11791 | 11791 | 11791 | 11791 | 8522 | 0
2017-07-24 22:00 | 11777 | 11777 | 11777 | 11777 | 11777 | 8522 | 0
2017-07-24 22:10 | 11771 | 11771 | 11771 | 11771 | 11771 | 8522 | 0
2017-07-24 22:20 | 11768 | 11768 | 11768 | 11768 | 11768 | 8522 | 0
2017-07-24 22:30 | 11768 | 11768 | 11768 | 11768 | 11768 | 8522 | 0
2017-07-24 22:40 | 11771 | 11771 | 11771 | 11771 | 11771 | 8522 | 0
2017-07-24 22:50 | 11771 | 11771 | 11771 | 11771 | 11771 | 8522 | 0
2017-07-24 23:00 | 11767 | 11767 | 11767 | 11767 | 11767 | 8522 | 0
2017-07-24 23:10 | 11774 | 11774 | 11774 | 11774 | 11774 | 8522 | 0
2017-07-24 23:20 | 11784 | 11784 | 11784 | 11784 | 11784 | 8522 | 0
2017-07-24 23:30 | 11779 | 11779 | 11779 | 11779 | 11779 | 8522 | 0
2017-07-24 23:40 | 11778 | 11778 | 11778 | 11778 | 11778 | 8522 | 0
2017-07-24 23:50 | 11761 | 11761 | 11761 | 11761 | 11761 | 8522 | 0
2017-07-25 | 11734 | 11734 | 11734 | 11734 | 11734 | 8522 | 0
2017-07-25 00:10 | 11694 | 11694 | 11694 | 11694 | 11694 | 8522 | 0
2017-07-25 00:20 | 11611 | 11611 | 11611 | 11611 | 11611 | 8522 | 0
2017-07-25 00:30 | 11544 | 11544 | 11544 | 11544 | 11544 | 8522 | 0
2017-07-25 00:40 | 11531 | 11531 | 11531 | 11531 | 11531 | 8522 | 0
2017-07-25 00:50 | 11543 | 11543 | 11543 | 11543 | 11543 | 8522 | 0
2017-07-25 01:00 | 11508 | 11508 | 11508 | 11508 | 11508 | 8522 | 0
2017-07-25 01:10 | 11408 | 11408 | 11408 | 11408 | 11408 | 8522 | 0
2017-07-25 01:20 | 11297 | 11297 | 11297 | 11297 | 11297 | 8522 | 0
2017-07-25 01:30 | 11206 | 11206 | 11206 | 11206 | 11206 | 8522 | 0
2017-07-25 01:40 | 11134 | 11134 | 11134 | 11134 | 11134 | 8522 | 0
2017-07-25 01:50 | 11110 | 11110 | 11110 | 11110 | 11110 | 8522 | 0
2017-07-25 02:00 | 11152 | 11152 | 11152 | 11152 | 11152 | 8522 | 0
2017-07-25 02:10 | 11210 | 11210 | 11210 | 11210 | 11210 | 8522 | 0
2017-07-25 02:20 | 11165 | 11165 | 11165 | 11165 | 11165 | 8522 | 0
2017-07-25 02:30 | 11097 | 11097 | 11097 | 11097 | 11097 | 8522 | 0
2017-07-25 02:40 | 11051 | 11051 | 11051 | 11051 | 11051 | 8522 | 0
2017-07-25 02:50 | 11044 | 11044 | 11044 | 11044 | 11044 | 8522 | 0
2017-07-25 03:00 | 11004 | 11004 | 11004 | 11004 | 11004 | 8522 | 0
2017-07-25 03:10 | 10954 | 10954 | 10954 | 10954 | 10954 | 8522 | 0
2017-07-25 03:20 | 10934 | 10934 | 10934 | 10934 | 10934 | 8522 | 0
2017-07-25 03:30 | 10936 | 10936 | 10936 | 10936 | 10936 | 8522 | 0
2017-07-25 03:40 | 10981 | 10981 | 10981 | 10981 | 10981 | 8522 | 0
2017-07-25 03:50 | 11027 | 11027 | 11027 | 11027 | 11027 | 8522 | 0
2017-07-25 04:00 | 10987 | 10987 | 10987 | 10987 | 10987 | 8522 | 0
2017-07-25 04:10 | 10921 | 10921 | 10921 | 10921 | 10921 | 8522 | 0
2017-07-25 04:20 | 10854 | 10854 | 10854 | 10854 | 10854 | 8522 | 0
2017-07-25 04:30 | 10804 | 10804 | 10804 | 10804 | 10804 | 8522 | 0
2017-07-25 04:40 | 10770 | 10770 | 10770 | 10770 | 10770 | 8522 | 0
2017-07-25 04:50 | 10725 | 10725 | 10725 | 10725 | 10725 | 8522 | 0
2017-07-25 05:00 | 10719 | 10719 | 10719 | 10719 | 10719 | 8522 | 0
2017-07-25 05:10 | 10767 | 10767 | 10767 | 10767 | 10767 | 8522 | 0
2017-07-25 05:20 | 10810 | 10810 | 10810 | 10810 | 10810 | 8522 | 0
2017-07-25 05:30 | 10879 | 10879 | 10879 | 10879 | 10879 | 8522 | 0
2017-07-25 05:40 | 10941 | 10941 | 10941 | 10941 | 10941 | 8522 | 0
2017-07-25 05:50 | 11000 | 11000 | 11000 | 11000 | 11000 | 8522 | 0
2017-07-25 06:00 | 11059 | 11059 | 11059 | 11059 | 11059 | 8522 | 0
2017-07-25 06:10 | 11115 | 11115 | 11115 | 11115 | 11115 | 8522 | 0
2017-07-25 06:20 | 11171 | 11171 | 11171 | 11171 | 11171 | 8522 | 0
2017-07-25 06:30 | 11226 | 11226 | 11226 | 11226 | 11226 | 8522 | 0
2017-07-25 06:40 | 11281 | 11281 | 11281 | 11281 | 11281 | 8522 | 0
2017-07-25 06:50 | 11335 | 11335 | 11335 | 11335 | 11335 | 8522 | 0
2017-07-25 07:00 | 11387 | 11387 | 11387 | 11387 | 11387 | 8522 | 0
2017-07-25 07:10 | 11439 | 11439 | 11439 | 11439 | 11439 | 8522 | 0
2017-07-25 07:20 | 11490 | 11490 | 11490 | 11490 | 11490 | 8522 | 0
2017-07-25 07:30 | 11539 | 11539 | 11539 | 11539 | 11539 | 8522 | 0
2017-07-25 07:40 | 11588 | 11588 | 11588 | 11588 | 11588 | 8522 | 0
2017-07-25 07:50 | 11635 | 11635 | 11635 | 11635 | 11635 | 8522 | 0
2017-07-25 08:00 | 11682 | 11682 | 11682 | 11682 | 11682 | 8522 | 0
2017-07-25 08:10 | 11727 | 11727 | 11727 | 11727 | 11727 | 8522 | 0
2017-07-25 08:20 | 11769 | 11769 | 11769 | 11769 | 11769 | 8522 | 0
2017-07-25 08:30 | 11810 | 11810 | 11810 | 11810 | 11810 | 8522 | 0
2017-07-25 08:40 | 11850 | 11850 | 11850 | 11850 | 11850 | 8522 | 0
2017-07-25 08:50 | 11889 | 11889 | 11889 | 11889 | 11889 | 8522 | 0
2017-07-25 09:00 | 11928 | 11928 | 11928 | 11928 | 11928 | 8522 | 0
2017-07-25 09:10 | 11966 | 11966 | 11966 | 11966 | 11966 | 8522 | 0
2017-07-25 09:20 | 12005 | 12005 | 12005 | 12005 | 12005 | 8522 | 0
2017-07-25 09:30 | 12042 | 12042 | 12042 | 12042 | 12042 | 8523 | 0
2017-07-25 09:40 | 12077 | 12077 | 12077 | 12077 | 12077 | 8522 | 0
2017-07-25 09:50 | 12112 | 12112 | 12112 | 12112 | 12112 | 8522 | 0
2017-07-25 10:00 | 12146 | 12146 | 12146 | 12146 | 12146 | 8522 | 0
2017-07-25 10:10 | 12180 | 12180 | 12180 | 12180 | 12180 | 8523 | 0
2017-07-25 10:20 | 12214 | 12214 | 12214 | 12214 | 12214 | 8522 | 0
2017-07-25 10:30 | 12248 | 12248 | 12248 | 12248 | 12248 | 8522 | 0
2017-07-25 10:40 | 12281 | 12281 | 12281 | 12281 | 12281 | 8522 | 0
2017-07-25 10:50 | 12314 | 12314 | 12314 | 12314 | 12314 | 8523 | 0
2017-07-25 11:00 | 12348 | 12348 | 12348 | 12348 | 12348 | 8522 | 0
2017-07-25 11:10 | 12381 | 12381 | 12381 | 12381 | 12381 | 8523 | 0
2017-07-25 11:20 | 12413 | 12413 | 12413 | 12413 | 12413 | 8522 | 0
2017-07-25 11:30 | 12445 | 12445 | 12445 | 12445 | 12445 | 8523 | 0
2017-07-25 11:40 | 12477 | 12477 | 12477 | 12477 | 12477 | 8522 | 0
2017-07-25 11:50 | 12509 | 12509 | 12509 | 12509 | 12509 | 8522 | 0
2017-07-25 12:00 | 12540 | 12540 | 12540 | 12540 | 12540 | 8522 | 0
2017-07-25 12:10 | 12571 | 12571 | 12571 | 12571 | 12571 | 8522 | 0
2017-07-25 12:20 | 12601 | 12601 | 12601 | 12601 | 12601 | 8522 | 0
2017-07-25 12:30 | 12631 | 12631 | 12631 | 12631 | 12631 | 8523 | 0
2017-07-25 12:40 | 12659 | 12659 | 12659 | 12659 | 12659 | 8522 | 0
2017-07-25 12:50 | 12687 | 12687 | 12687 | 12687 | 12687 | 8522 | 0
2017-07-25 13:00 | 12714 | 12714 | 12714 | 12714 | 12714 | 8522 | 0
2017-07-25 13:10 | 12740 | 12740 | 12740 | 12740 | 12740 | 8523 | 0
2017-07-25 13:20 | 12765 | 12765 | 12765 | 12765 | 12765 | 8522 | 0
2017-07-25 13:30 | 12788 | 12788 | 12788 | 12788 | 12788 | 8522 | 0
2017-07-25 13:40 | 12809 | 12809 | 12809 | 12809 | 12809 | 8522 | 0
2017-07-25 13:50 | 12829 | 12829 | 12829 | 12829 | 12829 | 8523 | 0
2017-07-25 14:00 | 12846 | 12846 | 12846 | 12846 | 12846 | 8522 | 0
2017-07-25 14:10 | 12861 | 12861 | 12861 | 12861 | 12861 | 8522 | 0
2017-07-25 14:20 | 12874 | 12874 | 12874 | 12874 | 12874 | 8523 | 0
2017-07-25 14:30 | 12886 | 12886 | 12886 | 12886 | 12886 | 8523 | 0
2017-07-25 14:40 | 12898 | 12898 | 12898 | 12898 | 12898 | 8522 | 0
2017-07-25 14:50 | 12908 | 12908 | 12908 | 12908 | 12908 | 8523 | 0
2017-07-25 15:00 | 12910 | 12910 | 12910 | 12910 | 12910 | 8522 | 0
2017-07-25 15:10 | 12907 | 12907 | 12907 | 12907 | 12907 | 8523 | 0
2017-07-25 15:20 | 12904 | 12904 | 12904 | 12904 | 12904 | 8523 | 0
2017-07-25 15:30 | 12898 | 12898 | 12898 | 12898 | 12898 | 8522 | 0
2017-07-25 15:40 | 12889 | 12889 | 12889 | 12889 | 12889 | 8522 | 0
2017-07-25 15:50 | 12874 | 12874 | 12874 | 12874 | 12874 | 8522 | 0
2017-07-25 16:00 | 12856 | 12856 | 12856 | 12856 | 12856 | 8522 | 0
2017-07-25 16:10 | 12838 | 12838 | 12838 | 12838 | 12838 | 8522 | 0
2017-07-25 16:20 | 12813 | 12813 | 12813 | 12813 | 12813 | 8522 | 0
2017-07-25 16:30 | 12797 | 12797 | 12797 | 12797 | 12797 | 8523 | 0
2017-07-25 16:40 | 12784 | 12784 | 12784 | 12784 | 12784 | 8523 | 0
2017-07-25 16:50 | 12771 | 12771 | 12771 | 12771 | 12771 | 8523 | 0
2017-07-25 17:00 | 12761 | 12761 | 12761 | 12761 | 12761 | 8523 | 0
2017-07-25 17:10 | 12749 | 12749 | 12749 | 12749 | 12749 | 8523 | 0
2017-07-25 17:20 | 12738 | 12738 | 12738 | 12738 | 12738 | 8522 | 0
2017-07-25 17:30 | 12728 | 12728 | 12728 | 12728 | 12728 | 8523 | 0
2017-07-25 17:40 | 12718 | 12718 | 12718 | 12718 | 12718 | 8522 | 0
2017-07-25 17:50 | 12710 | 12710 | 12710 | 12710 | 12710 | 8522 | 0
2017-07-25 18:00 | 12703 | 12703 | 12703 | 12703 | 12703 | 8522 | 0
2017-07-25 18:10 | 12697 | 12697 | 12697 | 12697 | 12697 | 8522 | 0
2017-07-25 18:20 | 12692 | 12692 | 12692 | 12692 | 12692 | 8522 | 0
2017-07-25 18:30 | 12678 | 12678 | 12678 | 12678 | 12678 | 8522 | 0
2017-07-25 18:40 | 12646 | 12646 | 12646 | 12646 | 12646 | 8522 | 0
2017-07-25 18:50 | 12633 | 12633 | 12633 | 12633 | 12633 | 8522 | 0
2017-07-25 19:00 | 12625 | 12625 | 12625 | 12625 | 12625 | 8522 | 0
2017-07-25 19:10 | 12614 | 12614 | 12614 | 12614 | 12614 | 8523 | 0
2017-07-25 19:20 | 12605 | 12605 | 12605 | 12605 | 12605 | 8522 | 0
2017-07-25 19:30 | 12596 | 12596 | 12596 | 12596 | 12596 | 8522 | 0
2017-07-25 19:40 | 12586 | 12586 | 12586 | 12586 | 12586 | 8522 | 0
2017-07-25 19:50 | 12580 | 12580 | 12580 | 12580 | 12580 | 8522 | 0
2017-07-25 20:00 | 12579 | 12579 | 12579 | 12579 | 12579 | 8523 | 0
2017-07-25 20:10 | 12568 | 12568 | 12568 | 12568 | 12568 | 8522 | 0
2017-07-25 20:20 | 12541 | 12541 | 12541 | 12541 | 12541 | 8522 | 0
2017-07-25 20:30 | 12496 | 12496 | 12496 | 12496 | 12496 | 8522 | 0
2017-07-25 20:40 | 12448 | 12448 | 12448 | 12448 | 12448 | 8522 | 0
2017-07-25 20:50 | 12416 | 12416 | 12416 | 12416 | 12416 | 8522 | 0
2017-07-25 21:00 | 12390 | 12390 | 12390 | 12390 | 12390 | 8522 | 0
2017-07-25 21:10 | 12366 | 12366 | 12366 | 12366 | 12366 | 8523 | 0
2017-07-25 21:20 | 12342 | 12342 | 12342 | 12342 | 12342 | 8523 | 0
2017-07-25 21:30 | 12317 | 12317 | 12317 | 12317 | 12317 | 8522 | 0
2017-07-25 21:40 | 12259 | 12259 | 12259 | 12259 | 12259 | 8522 | 0
2017-07-25 21:50 | 12204 | 12204 | 12204 | 12204 | 12204 | 8522 | 0
2017-07-25 22:00 | 12189 | 12189 | 12189 | 12189 | 12189 | 8522 | 0
2017-07-25 22:10 | 12161 | 12161 | 12161 | 12161 | 12161 | 8523 | 0
2017-07-25 22:20 | 12105 | 12105 | 12105 | 12105 | 12105 | 8522 | 0
2017-07-25 22:30 | 12029 | 12029 | 12029 | 12029 | 12029 | 8522 | 0
2017-07-25 22:40 | 11952 | 11952 | 11952 | 11952 | 11952 | 8522 | 0
2017-07-25 22:50 | 11896 | 11896 | 11896 | 11896 | 11896 | 8522 | 0
2017-07-25 23:00 | 11875 | 11875 | 11875 | 11875 | 11875 | 8522 | 0
2017-07-25 23:10 | 11879 | 11879 | 11879 | 11879 | 11879 | 8522 | 0
2017-07-25 23:20 | 11882 | 11882 | 11882 | 11882 | 11882 | 8522 | 0
2017-07-25 23:30 | 11887 | 11887 | 11887 | 11887 | 11887 | 8522 | 0
2017-07-25 23:40 | 11894 | 11894 | 11894 | 11894 | 11894 | 8522 | 0
2017-07-25 23:50 | 11901 | 11901 | 11901 | 11901 | 11901 | 8522 | 0
2017-07-26 | 11907 | 11907 | 11907 | 11907 | 11907 | 8522 | 0
2017-07-26 00:10 | 11914 | 11914 | 11914 | 11914 | 11914 | 8522 | 0
2017-07-26 00:20 | 11920 | 11920 | 11920 | 11920 | 11920 | 8523 | 0
2017-07-26 00:30 | 11925 | 11925 | 11925 | 11925 | 11925 | 8522 | 0
2017-07-26 00:40 | 11930 | 11930 | 11930 | 11930 | 11930 | 8522 | 0
2017-07-26 00:50 | 11935 | 11935 | 11935 | 11935 | 11935 | 8522 | 0
2017-07-26 01:00 | 11942 | 11942 | 11942 | 11942 | 11942 | 8522 | 0
2017-07-26 01:10 | 11950 | 11950 | 11950 | 11950 | 11950 | 8523 | 0
2017-07-26 01:20 | 11961 | 11961 | 11961 | 11961 | 11961 | 8522 | 0
2017-07-26 01:30 | 11970 | 11970 | 11970 | 11970 | 11970 | 8522 | 0
2017-07-26 01:40 | 11979 | 11979 | 11979 | 11979 | 11979 | 8522 | 0
2017-07-26 01:50 | 11986 | 11986 | 11986 | 11986 | 11986 | 8522 | 0
2017-07-26 02:00 | 11991 | 11991 | 11991 | 11991 | 11991 | 8523 | 0
2017-07-26 02:10 | 11996 | 11996 | 11996 | 11996 | 11996 | 8522 | 0
2017-07-26 02:20 | 12000 | 12000 | 12000 | 12000 | 12000 | 8523 | 0
2017-07-26 02:30 | 12003 | 12003 | 12003 | 12003 | 12003 | 8522 | 0
2017-07-26 02:40 | 12007 | 12007 | 12007 | 12007 | 12007 | 8522 | 0
2017-07-26 02:50 | 12009 | 12009 | 12009 | 12009 | 12009 | 8522 | 0
2017-07-26 03:00 | 12008 | 12008 | 12008 | 12008 | 12008 | 8522 | 0
2017-07-26 03:10 | 12005 | 12005 | 12005 | 12005 | 12005 | 8522 | 0
2017-07-26 03:20 | 12001 | 12001 | 12001 | 12001 | 12001 | 8522 | 0
2017-07-26 03:30 | 11996 | 11996 | 11996 | 11996 | 11996 | 8522 | 0
2017-07-26 03:40 | 11986 | 11986 | 11986 | 11986 | 11986 | 8522 | 0
2017-07-26 03:50 | 11974 | 11974 | 11974 | 11974 | 11974 | 8522 | 0
2017-07-26 04:00 | 11965 | 11965 | 11965 | 11965 | 11965 | 8522 | 0
2017-07-26 04:10 | 11962 | 11962 | 11962 | 11962 | 11962 | 8522 | 0
2017-07-26 04:20 | 11962 | 11962 | 11962 | 11962 | 11962 | 8522 | 0
2017-07-26 04:30 | 11967 | 11967 | 11967 | 11967 | 11967 | 8522 | 0
2017-07-26 04:40 | 11977 | 11977 | 11977 | 11977 | 11977 | 8522 | 0
2017-07-26 04:50 | 11990 | 11990 | 11990 | 11990 | 11990 | 8522 | 0
2017-07-26 05:00 | 11995 | 11995 | 11995 | 11995 | 11995 | 8522 | 0
2017-07-26 05:10 | 12011 | 12011 | 12011 | 12011 | 12011 | 8522 | 0
2017-07-26 05:20 | 12024 | 12024 | 12024 | 12024 | 12024 | 8522 | 0
2017-07-26 05:30 | 12035 | 12035 | 12035 | 12035 | 12035 | 8522 | 0
2017-07-26 05:40 | 12046 | 12046 | 12046 | 12046 | 12046 | 8522 | 0
2017-07-26 05:50 | 12052 | 12052 | 12052 | 12052 | 12052 | 8522 | 0
2017-07-26 06:00 | 12042 | 12042 | 12042 | 12042 | 12042 | 8522 | 0
2017-07-26 06:10 | 12039 | 12039 | 12039 | 12039 | 12039 | 8522 | 0
2017-07-26 06:20 | 12043 | 12043 | 12043 | 12043 | 12043 | 8522 | 0
2017-07-26 06:30 | 12050 | 12050 | 12050 | 12050 | 12050 | 8522 | 0
2017-07-26 06:40 | 12060 | 12060 | 12060 | 12060 | 12060 | 8522 | 0
2017-07-26 06:50 | 12074 | 12074 | 12074 | 12074 | 12074 | 8522 | 0
2017-07-26 07:00 | 12091 | 12091 | 12091 | 12091 | 12091 | 8522 | 0
2017-07-26 07:10 | 12110 | 12110 | 12110 | 12110 | 12110 | 8522 | 0
2017-07-26 07:20 | 12130 | 12130 | 12130 | 12130 | 12130 | 8522 | 0
2017-07-26 07:30 | 12154 | 12154 | 12154 | 12154 | 12154 | 8522 | 0
2017-07-26 07:40 | 12179 | 12179 | 12179 | 12179 | 12179 | 8522 | 0
2017-07-26 07:50 | 12207 | 12207 | 12207 | 12207 | 12207 | 8522 | 0
2017-07-26 08:00 | 12238 | 12238 | 12238 | 12238 | 12238 | 8523 | 0
2017-07-26 08:10 | 12270 | 12270 | 12270 | 12270 | 12270 | 8522 | 0
2017-07-26 08:20 | 12303 | 12303 | 12303 | 12303 | 12303 | 8522 | 0
2017-07-26 08:30 | 12336 | 12336 | 12336 | 12336 | 12336 | 8522 | 0
2017-07-26 08:40 | 12370 | 12370 | 12370 | 12370 | 12370 | 8522 | 0
2017-07-26 08:50 | 12404 | 12404 | 12404 | 12404 | 12404 | 8522 | 0
2017-07-26 09:00 | 12436 | 12436 | 12436 | 12436 | 12436 | 8522 | 0
2017-07-26 09:10 | 12468 | 12468 | 12468 | 12468 | 12468 | 8522 | 0
2017-07-26 09:20 | 12500 | 12500 | 12500 | 12500 | 12500 | 8522 | 0
2017-07-26 09:30 | 12532 | 12532 | 12532 | 12532 | 12532 | 8523 | 0
2017-07-26 09:40 | 12563 | 12563 | 12563 | 12563 | 12563 | 8523 | 0
2017-07-26 09:50 | 12594 | 12594 | 12594 | 12594 | 12594 | 8522 | 0
2017-07-26 10:00 | 12624 | 12624 | 12624 | 12624 | 12624 | 8522 | 0
2017-07-26 10:10 | 12655 | 12655 | 12655 | 12655 | 12655 | 8522 | 0
2017-07-26 10:20 | 12685 | 12685 | 12685 | 12685 | 12685 | 8522 | 0
2017-07-26 10:30 | 12716 | 12716 | 12716 | 12716 | 12716 | 8522 | 0
2017-07-26 10:40 | 12747 | 12747 | 12747 | 12747 | 12747 | 8522 | 0
2017-07-26 10:50 | 12778 | 12778 | 12778 | 12778 | 12778 | 8522 | 0
2017-07-26 11:00 | 12808 | 12808 | 12808 | 12808 | 12808 | 8522 | 0
2017-07-26 11:10 | 12837 | 12837 | 12837 | 12837 | 12837 | 8522 | 0
2017-07-26 11:20 | 12865 | 12865 | 12865 | 12865 | 12865 | 8522 | 0
2017-07-26 11:30 | 12892 | 12892 | 12892 | 12892 | 12892 | 8522 | 0
2017-07-26 11:40 | 12916 | 12916 | 12916 | 12916 | 12916 | 8522 | 0
2017-07-26 11:50 | 12938 | 12938 | 12938 | 12938 | 12938 | 8522 | 0
2017-07-26 12:00 | 12958 | 12958 | 12958 | 12958 | 12958 | 8522 | 0
2017-07-26 12:10 | 12978 | 12978 | 12978 | 12978 | 12978 | 8522 | 0
2017-07-26 12:20 | 12999 | 12999 | 12999 | 12999 | 12999 | 8522 | 0
2017-07-26 12:30 | 13019 | 13019 | 13019 | 13019 | 13019 | 8522 | 0
2017-07-26 12:40 | 13039 | 13039 | 13039 | 13039 | 13039 | 8522 | 0
2017-07-26 12:50 | 13059 | 13059 | 13059 | 13059 | 13059 | 8522 | 0
2017-07-26 13:00 | 13078 | 13078 | 13078 | 13078 | 13078 | 8522 | 0
2017-07-26 13:10 | 13098 | 13098 | 13098 | 13098 | 13098 | 8522 | 0
2017-07-26 13:20 | 13118 | 13118 | 13118 | 13118 | 13118 | 8522 | 0
2017-07-26 13:30 | 13138 | 13138 | 13138 | 13138 | 13138 | 8522 | 0
2017-07-26 13:40 | 13150 | 13150 | 13150 | 13150 | 13150 | 8522 | 0
2017-07-26 13:50 | 13159 | 13159 | 13159 | 13159 | 13159 | 8522 | 0
2017-07-26 14:00 | 13172 | 13172 | 13172 | 13172 | 13172 | 8522 | 0
2017-07-26 14:10 | 13188 | 13188 | 13188 | 13188 | 13188 | 8522 | 0
2017-07-26 14:20 | 13201 | 13201 | 13201 | 13201 | 13201 | 8522 | 0
2017-07-26 14:30 | 13213 | 13213 | 13213 | 13213 | 13213 | 8522 | 0
2017-07-26 14:40 | 13222 | 13222 | 13222 | 13222 | 13222 | 8522 | 0
2017-07-26 14:50 | 13225 | 13225 | 13225 | 13225 | 13225 | 8522 | 0
2017-07-26 15:00 | 13219 | 13219 | 13219 | 13219 | 13219 | 8522 | 0
2017-07-26 15:10 | 13212 | 13212 | 13212 | 13212 | 13212 | 8522 | 0
2017-07-26 15:20 | 13207 | 13207 | 13207 | 13207 | 13207 | 8522 | 0
2017-07-26 15:30 | 13203 | 13203 | 13203 | 13203 | 13203 | 8522 | 0
2017-07-26 15:40 | 13194 | 13194 | 13194 | 13194 | 13194 | 8522 | 0
2017-07-26 15:50 | 13180 | 13180 | 13180 | 13180 | 13180 | 8522 | 0
2017-07-26 16:00 | 13162 | 13162 | 13162 | 13162 | 13162 | 8522 | 0
2017-07-26 16:10 | 13143 | 13143 | 13143 | 13143 | 13143 | 8522 | 0
2017-07-26 16:20 | 13122 | 13122 | 13122 | 13122 | 13122 | 8522 | 0
2017-07-26 16:30 | 13100 | 13100 | 13100 | 13100 | 13100 | 8523 | 0
2017-07-26 16:40 | 13074 | 13074 | 13074 | 13074 | 13074 | 8522 | 0
2017-07-26 16:50 | 13046 | 13046 | 13046 | 13046 | 13046 | 8522 | 0
2017-07-26 17:00 | 13016 | 13016 | 13016 | 13016 | 13016 | 8522 | 0
2017-07-26 17:10 | 12984 | 12984 | 12984 | 12984 | 12984 | 8522 | 0
2017-07-26 17:20 | 12953 | 12953 | 12953 | 12953 | 12953 | 8522 | 0
2017-07-26 17:30 | 12924 | 12924 | 12924 | 12924 | 12924 | 8522 | 0
2017-07-26 17:40 | 12897 | 12897 | 12897 | 12897 | 12897 | 8522 | 0
2017-07-26 17:50 | 12872 | 12872 | 12872 | 12872 | 12872 | 8522 | 0
2017-07-26 18:00 | 12843 | 12843 | 12843 | 12843 | 12843 | 8522 | 0
2017-07-26 18:10 | 12811 | 12811 | 12811 | 12811 | 12811 | 8522 | 0
2017-07-26 18:20 | 12774 | 12774 | 12774 | 12774 | 12774 | 8522 | 0
2017-07-26 18:30 | 12741 | 12741 | 12741 | 12741 | 12741 | 8523 | 0
2017-07-26 18:40 | 12711 | 12711 | 12711 | 12711 | 12711 | 8522 | 0
2017-07-26 18:50 | 12683 | 12683 | 12683 | 12683 | 12683 | 8522 | 0
2017-07-26 19:00 | 12658 | 12658 | 12658 | 12658 | 12658 | 8522 | 0
2017-07-26 19:10 | 12633 | 12633 | 12633 | 12633 | 12633 | 8523 | 0
2017-07-26 19:20 | 12607 | 12607 | 12607 | 12607 | 12607 | 8522 | 0
2017-07-26 19:30 | 12575 | 12575 | 12575 | 12575 | 12575 | 8522 | 0
2017-07-26 19:40 | 12542 | 12542 | 12542 | 12542 | 12542 | 8522 | 0
2017-07-26 19:50 | 12492 | 12492 | 12492 | 12492 | 12492 | 8522 | 0
2017-07-26 20:00 | 12429 | 12429 | 12429 | 12429 | 12429 | 8522 | 0
2017-07-26 20:10 | 12378 | 12378 | 12378 | 12378 | 12378 | 8522 | 0
2017-07-26 20:20 | 12337 | 12337 | 12337 | 12337 | 12337 | 8522 | 0
2017-07-26 20:30 | 12298 | 12298 | 12298 | 12298 | 12298 | 8522 | 0
2017-07-26 20:40 | 12262 | 12262 | 12262 | 12262 | 12262 | 8522 | 0
2017-07-26 20:50 | 12214 | 12214 | 12214 | 12214 | 12214 | 8522 | 0
2017-07-26 21:00 | 12130 | 12130 | 12130 | 12130 | 12130 | 8522 | 0
2017-07-26 21:10 | 12039 | 12039 | 12039 | 12039 | 12039 | 8522 | 0
2017-07-26 21:20 | 11874 | 11874 | 11874 | 11874 | 11874 | 8522 | 0
2017-07-26 21:30 | 11716 | 11716 | 11716 | 11716 | 11716 | 8522 | 0
2017-07-26 21:40 | 11568 | 11568 | 11568 | 11568 | 11568 | 8522 | 0
2017-07-26 21:50 | 11429 | 11429 | 11429 | 11429 | 11429 | 8522 | 0
2017-07-26 22:00 | 11304 | 11304 | 11304 | 11304 | 11304 | 8522 | 0
2017-07-26 22:10 | 11212 | 11212 | 11212 | 11212 | 11212 | 8522 | 0
2017-07-26 22:20 | 11169 | 11169 | 11169 | 11169 | 11169 | 8522 | 0
2017-07-26 22:30 | 11168 | 11168 | 11168 | 11168 | 11168 | 8522 | 0
2017-07-26 22:40 | 11182 | 11182 | 11182 | 11182 | 11182 | 8522 | 0
2017-07-26 22:50 | 11201 | 11201 | 11201 | 11201 | 11201 | 8522 | 0
2017-07-26 23:00 | 11212 | 11212 | 11212 | 11212 | 11212 | 8522 | 0
2017-07-26 23:10 | 11200 | 11200 | 11200 | 11200 | 11200 | 8522 | 0
2017-07-26 23:20 | 11189 | 11189 | 11189 | 11189 | 11189 | 8522 | 0
2017-07-26 23:30 | 11180 | 11180 | 11180 | 11180 | 11180 | 8522 | 0
2017-07-26 23:40 | 11172 | 11172 | 11172 | 11172 | 11172 | 8522 | 0
2017-07-26 23:50 | 11167 | 11167 | 11167 | 11167 | 11167 | 8522 | 0
2017-07-27 | 11163 | 11163 | 11163 | 11163 | 11163 | 8522 | 0
2017-07-27 00:10 | 11148 | 11148 | 11148 | 11148 | 11148 | 8522 | 0
2017-07-27 00:20 | 11132 | 11132 | 11132 | 11132 | 11132 | 8522 | 0
2017-07-27 00:30 | 11118 | 11118 | 11118 | 11118 | 11118 | 8522 | 0
2017-07-27 00:40 | 11102 | 11102 | 11102 | 11102 | 11102 | 8522 | 0
2017-07-27 00:50 | 11079 | 11079 | 11079 | 11079 | 11079 | 8522 | 0
2017-07-27 01:00 | 11016 | 11016 | 11016 | 11016 | 11016 | 8522 | 0
2017-07-27 01:10 | 10951 | 10951 | 10951 | 10951 | 10951 | 8522 | 0
2017-07-27 01:20 | 10887 | 10887 | 10887 | 10887 | 10887 | 8522 | 0
2017-07-27 01:30 | 10844 | 10844 | 10844 | 10844 | 10844 | 8522 | 0
2017-07-27 01:40 | 10808 | 10808 | 10808 | 10808 | 10808 | 8522 | 0
2017-07-27 01:50 | 10775 | 10775 | 10775 | 10775 | 10775 | 8522 | 0
2017-07-27 02:00 | 10740 | 10740 | 10740 | 10740 | 10740 | 8522 | 0
2017-07-27 02:10 | 10655 | 10655 | 10655 | 10655 | 10655 | 8522 | 0
2017-07-27 02:20 | 10528 | 10528 | 10528 | 10528 | 10528 | 8522 | 0
2017-07-27 02:30 | 10398 | 10398 | 10398 | 10398 | 10398 | 8522 | 0
2017-07-27 02:40 | 10270 | 10270 | 10270 | 10270 | 10270 | 8522 | 0
2017-07-27 02:50 | 10179 | 10179 | 10179 | 10179 | 10179 | 8522 | 0
2017-07-27 03:00 | 10083 | 10083 | 10083 | 10083 | 10083 | 8522 | 0
2017-07-27 03:10 | 10038 | 10038 | 10038 | 10038 | 10038 | 8522 | 0
2017-07-27 03:20 | 10055 | 10055 | 10055 | 10055 | 10055 | 8522 | 0
2017-07-27 03:30 | 10093 | 10093 | 10093 | 10093 | 10093 | 8522 | 0
2017-07-27 03:40 | 10135 | 10135 | 10135 | 10135 | 10135 | 8522 | 0
2017-07-27 03:50 | 10174 | 10174 | 10174 | 10174 | 10174 | 8522 | 0
2017-07-27 04:00 | 10209 | 10209 | 10209 | 10209 | 10209 | 8522 | 0
2017-07-27 04:10 | 10239 | 10239 | 10239 | 10239 | 10239 | 8522 | 0
2017-07-27 04:20 | 10265 | 10265 | 10265 | 10265 | 10265 | 8522 | 0
2017-07-27 04:30 | 10286 | 10286 | 10286 | 10286 | 10286 | 8522 | 0
2017-07-27 04:40 | 10305 | 10305 | 10305 | 10305 | 10305 | 8522 | 0
2017-07-27 04:50 | 10321 | 10321 | 10321 | 10321 | 10321 | 8522 | 0
2017-07-27 05:00 | 10332 | 10332 | 10332 | 10332 | 10332 | 8522 | 0
2017-07-27 05:10 | 10328 | 10328 | 10328 | 10328 | 10328 | 8522 | 0
2017-07-27 05:20 | 10303 | 10303 | 10303 | 10303 | 10303 | 8522 | 0
2017-07-27 05:30 | 10259 | 10259 | 10259 | 10259 | 10259 | 8522 | 0
2017-07-27 05:40 | 10183 | 10183 | 10183 | 10183 | 10183 | 8522 | 0
2017-07-27 05:50 | 10143 | 10143 | 10143 | 10143 | 10143 | 8522 | 0
2017-07-27 06:00 | 10133 | 10133 | 10133 | 10133 | 10133 | 8522 | 0
2017-07-27 06:10 | 10145 | 10145 | 10145 | 10145 | 10145 | 8522 | 0
2017-07-27 06:20 | 10171 | 10171 | 10171 | 10171 | 10171 | 8522 | 0
2017-07-27 06:30 | 10206 | 10206 | 10206 | 10206 | 10206 | 8522 | 0
2017-07-27 06:40 | 10249 | 10249 | 10249 | 10249 | 10249 | 8522 | 0
2017-07-27 06:50 | 10296 | 10296 | 10296 | 10296 | 10296 | 8522 | 0
2017-07-27 07:00 | 10347 | 10347 | 10347 | 10347 | 10347 | 8522 | 0
2017-07-27 07:10 | 10400 | 10400 | 10400 | 10400 | 10400 | 8522 | 0
2017-07-27 07:20 | 10454 | 10454 | 10454 | 10454 | 10454 | 8522 | 0
2017-07-27 07:30 | 10510 | 10510 | 10510 | 10510 | 10510 | 8522 | 0
2017-07-27 07:40 | 10567 | 10567 | 10567 | 10567 | 10567 | 8522 | 0
2017-07-27 07:50 | 10623 | 10623 | 10623 | 10623 | 10623 | 8522 | 0
2017-07-27 08:00 | 10679 | 10679 | 10679 | 10679 | 10679 | 8522 | 0
2017-07-27 08:10 | 10734 | 10734 | 10734 | 10734 | 10734 | 8522 | 0
2017-07-27 08:20 | 10786 | 10786 | 10786 | 10786 | 10786 | 8522 | 0
2017-07-27 08:30 | 10837 | 10837 | 10837 | 10837 | 10837 | 8522 | 0
2017-07-27 08:40 | 10886 | 10886 | 10886 | 10886 | 10886 | 8522 | 0
2017-07-27 08:50 | 10933 | 10933 | 10933 | 10933 | 10933 | 8522 | 0
2017-07-27 09:00 | 10978 | 10978 | 10978 | 10978 | 10978 | 8522 | 0
2017-07-27 09:10 | 11021 | 11021 | 11021 | 11021 | 11021 | 8522 | 0
2017-07-27 09:20 | 11063 | 11063 | 11063 | 11063 | 11063 | 8522 | 0
2017-07-27 09:30 | 11104 | 11104 | 11104 | 11104 | 11104 | 8522 | 0
2017-07-27 09:40 | 11144 | 11144 | 11144 | 11144 | 11144 | 8522 | 0
2017-07-27 09:50 | 11184 | 11184 | 11184 | 11184 | 11184 | 8522 | 0
2017-07-27 10:00 | 11222 | 11222 | 11222 | 11222 | 11222 | 8522 | 0
2017-07-27 10:10 | 11259 | 11259 | 11259 | 11259 | 11259 | 8522 | 0
2017-07-27 10:20 | 11294 | 11294 | 11294 | 11294 | 11294 | 8522 | 0
2017-07-27 10:30 | 11328 | 11328 | 11328 | 11328 | 11328 | 8522 | 0
2017-07-27 10:40 | 11361 | 11361 | 11361 | 11361 | 11361 | 8522 | 0
2017-07-27 10:50 | 11394 | 11394 | 11394 | 11394 | 11394 | 8522 | 0
2017-07-27 11:00 | 11427 | 11427 | 11427 | 11427 | 11427 | 8522 | 0
2017-07-27 11:10 | 11458 | 11458 | 11458 | 11458 | 11458 | 8522 | 0
2017-07-27 11:20 | 11489 | 11489 | 11489 | 11489 | 11489 | 8522 | 0
2017-07-27 11:30 | 11520 | 11520 | 11520 | 11520 | 11520 | 8522 | 0
2017-07-27 11:40 | 11550 | 11550 | 11550 | 11550 | 11550 | 8522 | 0
2017-07-27 11:50 | 11579 | 11579 | 11579 | 11579 | 11579 | 8522 | 0
2017-07-27 12:00 | 11608 | 11608 | 11608 | 11608 | 11608 | 8522 | 0
2017-07-27 12:10 | 11636 | 11636 | 11636 | 11636 | 11636 | 8522 | 0
2017-07-27 12:20 | 11661 | 11661 | 11661 | 11661 | 11661 | 8522 | 0
2017-07-27 12:30 | 11683 | 11683 | 11683 | 11683 | 11683 | 8523 | 0
2017-07-27 12:40 | 11704 | 11704 | 11704 | 11704 | 11704 | 8522 | 0
2017-07-27 12:50 | 11724 | 11724 | 11724 | 11724 | 11724 | 8522 | 0
2017-07-27 13:00 | 11743 | 11743 | 11743 | 11743 | 11743 | 8522 | 0
2017-07-27 13:10 | 11761 | 11761 | 11761 | 11761 | 11761 | 8522 | 0
2017-07-27 13:20 | 11778 | 11778 | 11778 | 11778 | 11778 | 8522 | 0
2017-07-27 13:30 | 11795 | 11795 | 11795 | 11795 | 11795 | 8522 | 0
2017-07-27 13:40 | 11811 | 11811 | 11811 | 11811 | 11811 | 8522 | 0
2017-07-27 13:50 | 11826 | 11826 | 11826 | 11826 | 11826 | 8522 | 0
2017-07-27 14:00 | 11842 | 11842 | 11842 | 11842 | 11842 | 8522 | 0
2017-07-27 14:10 | 11858 | 11858 | 11858 | 11858 | 11858 | 8522 | 0
2017-07-27 14:20 | 11874 | 11874 | 11874 | 11874 | 11874 | 8523 | 0
2017-07-27 14:30 | 11891 | 11891 | 11891 | 11891 | 11891 | 8523 | 0
2017-07-27 14:40 | 11907 | 11907 | 11907 | 11907 | 11907 | 8522 | 0
2017-07-27 14:50 | 11919 | 11919 | 11919 | 11919 | 11919 | 8522 | 0
2017-07-27 15:00 | 11926 | 11926 | 11926 | 11926 | 11926 | 8522 | 0
2017-07-27 15:10 | 11930 | 11930 | 11930 | 11930 | 11930 | 8522 | 0
2017-07-27 15:20 | 11933 | 11933 | 11933 | 11933 | 11933 | 8522 | 0
2017-07-27 15:30 | 11933 | 11933 | 11933 | 11933 | 11933 | 8522 | 0
2017-07-27 15:40 | 11930 | 11930 | 11930 | 11930 | 11930 | 8523 | 0
2017-07-27 15:50 | 11922 | 11922 | 11922 | 11922 | 11922 | 8522 | 0
2017-07-27 16:00 | 11916 | 11916 | 11916 | 11916 | 11916 | 8522 | 0
2017-07-27 16:10 | 11914 | 11914 | 11914 | 11914 | 11914 | 8522 | 0
2017-07-27 16:20 | 11910 | 11910 | 11910 | 11910 | 11910 | 8522 | 0
2017-07-27 16:30 | 11904 | 11904 | 11904 | 11904 | 11904 | 8522 | 0
2017-07-27 16:40 | 11890 | 11890 | 11890 | 11890 | 11890 | 8522 | 0
2017-07-27 16:50 | 11870 | 11870 | 11870 | 11870 | 11870 | 8522 | 0
2017-07-27 17:00 | 11855 | 11855 | 11855 | 11855 | 11855 | 8522 | 0
2017-07-27 17:10 | 11848 | 11848 | 11848 | 11848 | 11848 | 8522 | 0
2017-07-27 17:20 | 11845 | 11845 | 11845 | 11845 | 11845 | 8522 | 0
2017-07-27 17:30 | 11844 | 11844 | 11844 | 11844 | 11844 | 8523 | 0
2017-07-27 17:40 | 11840 | 11840 | 11840 | 11840 | 11840 | 8523 | 0
2017-07-27 17:50 | 11835 | 11835 | 11835 | 11835 | 11835 | 8522 | 0
2017-07-27 18:00 | 11829 | 11829 | 11829 | 11829 | 11829 | 8522 | 0
2017-07-27 18:10 | 11822 | 11822 | 11822 | 11822 | 11822 | 8523 | 0
2017-07-27 18:20 | 11814 | 11814 | 11814 | 11814 | 11814 | 8522 | 0
2017-07-27 18:30 | 11805 | 11805 | 11805 | 11805 | 11805 | 8522 | 0
2017-07-27 18:40 | 11795 | 11795 | 11795 | 11795 | 11795 | 8522 | 0
2017-07-27 18:50 | 11785 | 11785 | 11785 | 11785 | 11785 | 8522 | 0
2017-07-27 19:00 | 11776 | 11776 | 11776 | 11776 | 11776 | 8522 | 0
2017-07-27 19:10 | 11767 | 11767 | 11767 | 11767 | 11767 | 8522 | 0
2017-07-27 19:20 | 11755 | 11755 | 11755 | 11755 | 11755 | 8522 | 0
2017-07-27 19:30 | 11732 | 11732 | 11732 | 11732 | 11732 | 8522 | 0
2017-07-27 19:40 | 11700 | 11700 | 11700 | 11700 | 11700 | 8522 | 0
2017-07-27 19:50 | 11662 | 11662 | 11662 | 11662 | 11662 | 8522 | 0
2017-07-27 20:00 | 11616 | 11616 | 11616 | 11616 | 11616 | 8522 | 0
2017-07-27 20:10 | 11513 | 11513 | 11513 | 11513 | 11513 | 8522 | 0
2017-07-27 20:20 | 11387 | 11387 | 11387 | 11387 | 11387 | 8522 | 0
2017-07-27 20:30 | 11249 | 11249 | 11249 | 11249 | 11249 | 8522 | 0
2017-07-27 20:40 | 11112 | 11112 | 11112 | 11112 | 11112 | 8522 | 0
2017-07-27 20:50 | 10993 | 10993 | 10993 | 10993 | 10993 | 8522 | 0
2017-07-27 21:00 | 10947 | 10947 | 10947 | 10947 | 10947 | 8522 | 0
2017-07-27 21:10 | 10947 | 10947 | 10947 | 10947 | 10947 | 8522 | 0
2017-07-27 21:20 | 10969 | 10969 | 10969 | 10969 | 10969 | 8522 | 0
2017-07-27 21:30 | 11004 | 11004 | 11004 | 11004 | 11004 | 8522 | 0
2017-07-27 21:40 | 11044 | 11044 | 11044 | 11044 | 11044 | 8522 | 0
2017-07-27 21:50 | 11085 | 11085 | 11085 | 11085 | 11085 | 8522 | 0
2017-07-27 22:00 | 11127 | 11127 | 11127 | 11127 | 11127 | 8522 | 0
2017-07-27 22:10 | 11166 | 11166 | 11166 | 11166 | 11166 | 8522 | 0
2017-07-27 22:20 | 11200 | 11200 | 11200 | 11200 | 11200 | 8522 | 0
2017-07-27 22:30 | 11230 | 11230 | 11230 | 11230 | 11230 | 8522 | 0
2017-07-27 22:40 | 11257 | 11257 | 11257 | 11257 | 11257 | 8522 | 0
2017-07-27 22:50 | 11282 | 11282 | 11282 | 11282 | 11282 | 8522 | 0
2017-07-27 23:00 | 11303 | 11303 | 11303 | 11303 | 11303 | 8522 | 0
2017-07-27 23:10 | 11326 | 11326 | 11326 | 11326 | 11326 | 8522 | 0
2017-07-27 23:20 | 11343 | 11343 | 11343 | 11343 | 11343 | 8522 | 0
2017-07-27 23:30 | 11349 | 11349 | 11349 | 11349 | 11349 | 8522 | 0
2017-07-27 23:40 | 11331 | 11331 | 11331 | 11331 | 11331 | 8522 | 0
2017-07-27 23:50 | 11296 | 11296 | 11296 | 11296 | 11296 | 8522 | 0
2017-07-28 | 11185 | 11185 | 11185 | 11185 | 11185 | 8522 | 0
2017-07-28 00:10 | 11017 | 11017 | 11017 | 11017 | 11017 | 8522 | 0
2017-07-28 00:20 | 10862 | 10862 | 10862 | 10862 | 10862 | 8522 | 0
2017-07-28 00:30 | 10753 | 10753 | 10753 | 10753 | 10753 | 8522 | 0
2017-07-28 00:40 | 10650 | 10650 | 10650 | 10650 | 10650 | 8522 | 0
2017-07-28 00:50 | 10552 | 10552 | 10552 | 10552 | 10552 | 8522 | 0
2017-07-28 01:00 | 10488 | 10488 | 10488 | 10488 | 10488 | 8522 | 0
2017-07-28 01:10 | 10417 | 10417 | 10417 | 10417 | 10417 | 8522 | 0
2017-07-28 01:20 | 10393 | 10393 | 10393 | 10393 | 10393 | 8522 | 0
2017-07-28 01:30 | 10436 | 10436 | 10436 | 10436 | 10436 | 8522 | 0
2017-07-28 01:40 | 10502 | 10502 | 10502 | 10502 | 10502 | 8522 | 0
2017-07-28 01:50 | 10570 | 10570 | 10570 | 10570 | 10570 | 8522 | 0
2017-07-28 02:00 | 10631 | 10631 | 10631 | 10631 | 10631 | 8522 | 0
2017-07-28 02:10 | 10665 | 10665 | 10665 | 10665 | 10665 | 8522 | 0
2017-07-28 02:20 | 10697 | 10697 | 10697 | 10697 | 10697 | 8522 | 0
2017-07-28 02:30 | 10724 | 10724 | 10724 | 10724 | 10724 | 8522 | 0
2017-07-28 02:40 | 10744 | 10744 | 10744 | 10744 | 10744 | 8522 | 0
2017-07-28 02:50 | 10752 | 10752 | 10752 | 10752 | 10752 | 8522 | 0
2017-07-28 03:00 | 10756 | 10756 | 10756 | 10756 | 10756 | 8522 | 0
2017-07-28 03:10 | 10763 | 10763 | 10763 | 10763 | 10763 | 8522 | 0
2017-07-28 03:20 | 10771 | 10771 | 10771 | 10771 | 10771 | 8522 | 0
2017-07-28 03:30 | 10733 | 10733 | 10733 | 10733 | 10733 | 8522 | 0
2017-07-28 03:40 | 10697 | 10697 | 10697 | 10697 | 10697 | 8522 | 0
2017-07-28 03:50 | 10676 | 10676 | 10676 | 10676 | 10676 | 8522 | 0
2017-07-28 04:00 | 10658 | 10658 | 10658 | 10658 | 10658 | 8522 | 0
2017-07-28 04:10 | 10640 | 10640 | 10640 | 10640 | 10640 | 8522 | 0
2017-07-28 04:20 | 10621 | 10621 | 10621 | 10621 | 10621 | 8522 | 0
2017-07-28 04:30 | 10602 | 10602 | 10602 | 10602 | 10602 | 8522 | 0
2017-07-28 04:40 | 10583 | 10583 | 10583 | 10583 | 10583 | 8522 | 0
2017-07-28 04:50 | 10564 | 10564 | 10564 | 10564 | 10564 | 8522 | 0
2017-07-28 05:00 | 10544 | 10544 | 10544 | 10544 | 10544 | 8522 | 0
2017-07-28 05:10 | 10537 | 10537 | 10537 | 10537 | 10537 | 8522 | 0
2017-07-28 05:20 | 10545 | 10545 | 10545 | 10545 | 10545 | 8522 | 0
2017-07-28 05:30 | 10562 | 10562 | 10562 | 10562 | 10562 | 8522 | 0
2017-07-28 05:40 | 10582 | 10582 | 10582 | 10582 | 10582 | 8522 | 0
2017-07-28 05:50 | 10606 | 10606 | 10606 | 10606 | 10606 | 8522 | 0
2017-07-28 06:00 | 10632 | 10632 | 10632 | 10632 | 10632 | 8522 | 0
2017-07-28 06:10 | 10656 | 10656 | 10656 | 10656 | 10656 | 8522 | 0
2017-07-28 06:20 | 10680 | 10680 | 10680 | 10680 | 10680 | 8522 | 0
2017-07-28 06:30 | 10702 | 10702 | 10702 | 10702 | 10702 | 8522 | 0
2017-07-28 06:40 | 10724 | 10724 | 10724 | 10724 | 10724 | 8522 | 0
2017-07-28 06:50 | 10747 | 10747 | 10747 | 10747 | 10747 | 8522 | 0
2017-07-28 07:00 | 10771 | 10771 | 10771 | 10771 | 10771 | 8522 | 0
2017-07-28 07:10 | 10795 | 10795 | 10795 | 10795 | 10795 | 8522 | 0
2017-07-28 07:20 | 10819 | 10819 | 10819 | 10819 | 10819 | 8522 | 0
2017-07-28 07:30 | 10843 | 10843 | 10843 | 10843 | 10843 | 8522 | 0
2017-07-28 07:40 | 10865 | 10865 | 10865 | 10865 | 10865 | 8522 | 0
2017-07-28 07:50 | 10888 | 10888 | 10888 | 10888 | 10888 | 8522 | 0
2017-07-28 08:00 | 10910 | 10910 | 10910 | 10910 | 10910 | 8522 | 0
2017-07-28 08:10 | 10933 | 10933 | 10933 | 10933 | 10933 | 8522 | 0
2017-07-28 08:20 | 10957 | 10957 | 10957 | 10957 | 10957 | 8522 | 0
2017-07-28 08:30 | 10980 | 10980 | 10980 | 10980 | 10980 | 8522 | 0
2017-07-28 08:40 | 11002 | 11002 | 11002 | 11002 | 11002 | 8522 | 0
2017-07-28 08:50 | 11023 | 11023 | 11023 | 11023 | 11023 | 8522 | 0
2017-07-28 09:00 | 11043 | 11043 | 11043 | 11043 | 11043 | 8522 | 0
2017-07-28 09:10 | 11063 | 11063 | 11063 | 11063 | 11063 | 8522 | 0
2017-07-28 09:20 | 11082 | 11082 | 11082 | 11082 | 11082 | 8522 | 0
2017-07-28 09:30 | 11101 | 11101 | 11101 | 11101 | 11101 | 8522 | 0
2017-07-28 09:40 | 11119 | 11119 | 11119 | 11119 | 11119 | 8522 | 0
2017-07-28 09:50 | 11137 | 11137 | 11137 | 11137 | 11137 | 8522 | 0
2017-07-28 10:00 | 11157 | 11157 | 11157 | 11157 | 11157 | 8522 | 0
2017-07-28 10:10 | 11176 | 11176 | 11176 | 11176 | 11176 | 8522 | 0
2017-07-28 10:20 | 11196 | 11196 | 11196 | 11196 | 11196 | 8522 | 0
2017-07-28 10:30 | 11217 | 11217 | 11217 | 11217 | 11217 | 8522 | 0
2017-07-28 10:40 | 11239 | 11239 | 11239 | 11239 | 11239 | 8522 | 0
2017-07-28 10:50 | 11261 | 11261 | 11261 | 11261 | 11261 | 8522 | 0
2017-07-28 11:00 | 11282 | 11282 | 11282 | 11282 | 11282 | 8522 | 0
2017-07-28 11:10 | 11305 | 11305 | 11305 | 11305 | 11305 | 8522 | 0
2017-07-28 11:20 | 11328 | 11328 | 11328 | 11328 | 11328 | 8522 | 0
2017-07-28 11:30 | 11351 | 11351 | 11351 | 11351 | 11351 | 8522 | 0
2017-07-28 11:40 | 11374 | 11374 | 11374 | 11374 | 11374 | 8522 | 0
2017-07-28 11:50 | 11398 | 11398 | 11398 | 11398 | 11398 | 8522 | 0
2017-07-28 12:00 | 11421 | 11421 | 11421 | 11421 | 11421 | 8522 | 0
2017-07-28 12:10 | 11442 | 11442 | 11442 | 11442 | 11442 | 8522 | 0
2017-07-28 12:20 | 11462 | 11462 | 11462 | 11462 | 11462 | 8522 | 0
2017-07-28 12:30 | 11481 | 11481 | 11481 | 11481 | 11481 | 8522 | 0
2017-07-28 12:40 | 11501 | 11501 | 11501 | 11501 | 11501 | 8522 | 0
2017-07-28 12:50 | 11520 | 11520 | 11520 | 11520 | 11520 | 8522 | 0
2017-07-28 13:00 | 11538 | 11538 | 11538 | 11538 | 11538 | 8522 | 0
2017-07-28 13:10 | 11555 | 11555 | 11555 | 11555 | 11555 | 8522 | 0
2017-07-28 13:20 | 11572 | 11572 | 11572 | 11572 | 11572 | 8522 | 0
2017-07-28 13:30 | 11589 | 11589 | 11589 | 11589 | 11589 | 8522 | 0
2017-07-28 13:40 | 11606 | 11606 | 11606 | 11606 | 11606 | 8523 | 0
2017-07-28 13:50 | 11623 | 11623 | 11623 | 11623 | 11623 | 8522 | 0
2017-07-28 14:00 | 11638 | 11638 | 11638 | 11638 | 11638 | 8522 | 0
2017-07-28 14:10 | 11654 | 11654 | 11654 | 11654 | 11654 | 8522 | 0
2017-07-28 14:20 | 11669 | 11669 | 11669 | 11669 | 11669 | 8522 | 0
2017-07-28 14:30 | 11679 | 11679 | 11679 | 11679 | 11679 | 8522 | 0
2017-07-28 14:40 | 11688 | 11688 | 11688 | 11688 | 11688 | 8522 | 0
2017-07-28 14:50 | 11697 | 11697 | 11697 | 11697 | 11697 | 8522 | 0
2017-07-28 15:00 | 11704 | 11704 | 11704 | 11704 | 11704 | 8522 | 0
2017-07-28 15:10 | 11709 | 11709 | 11709 | 11709 | 11709 | 8522 | 0
2017-07-28 15:20 | 11713 | 11713 | 11713 | 11713 | 11713 | 8522 | 0
2017-07-28 15:30 | 11715 | 11715 | 11715 | 11715 | 11715 | 8522 | 0
2017-07-28 15:40 | 11713 | 11713 | 11713 | 11713 | 11713 | 8522 | 0
2017-07-28 15:50 | 11709 | 11709 | 11709 | 11709 | 11709 | 8522 | 0
2017-07-28 16:00 | 11704 | 11704 | 11704 | 11704 | 11704 | 8522 | 0
2017-07-28 16:10 | 11697 | 11697 | 11697 | 11697 | 11697 | 8522 | 0
2017-07-28 16:20 | 11691 | 11691 | 11691 | 11691 | 11691 | 8522 | 0
2017-07-28 16:30 | 11683 | 11683 | 11683 | 11683 | 11683 | 8522 | 0
2017-07-28 16:40 | 11647 | 11647 | 11647 | 11647 | 11647 | 8522 | 0
2017-07-28 16:50 | 11606 | 11606 | 11606 | 11606 | 11606 | 8522 | 0
2017-07-28 17:00 | 11566 | 11566 | 11566 | 11566 | 11566 | 8522 | 0
2017-07-28 17:10 | 11543 | 11543 | 11543 | 11543 | 11543 | 8522 | 0
2017-07-28 17:20 | 11531 | 11531 | 11531 | 11531 | 11531 | 8522 | 0
2017-07-28 17:30 | 11524 | 11524 | 11524 | 11524 | 11524 | 8522 | 0
2017-07-28 17:40 | 11519 | 11519 | 11519 | 11519 | 11519 | 8522 | 0
2017-07-28 17:50 | 11515 | 11515 | 11515 | 11515 | 11515 | 8522 | 0
2017-07-28 18:00 | 11511 | 11511 | 11511 | 11511 | 11511 | 8522 | 0
2017-07-28 18:10 | 11506 | 11506 | 11506 | 11506 | 11506 | 8522 | 0
2017-07-28 18:20 | 11501 | 11501 | 11501 | 11501 | 11501 | 8522 | 0
2017-07-28 18:30 | 11495 | 11495 | 11495 | 11495 | 11495 | 8522 | 0
2017-07-28 18:40 | 11489 | 11489 | 11489 | 11489 | 11489 | 8522 | 0
2017-07-28 18:50 | 11481 | 11481 | 11481 | 11481 | 11481 | 8522 | 0
2017-07-28 19:00 | 11472 | 11472 | 11472 | 11472 | 11472 | 8522 | 0
2017-07-28 19:10 | 11461 | 11461 | 11461 | 11461 | 11461 | 8522 | 0
2017-07-28 19:20 | 11452 | 11452 | 11452 | 11452 | 11452 | 8522 | 0
2017-07-28 19:30 | 11443 | 11443 | 11443 | 11443 | 11443 | 8522 | 0
2017-07-28 19:40 | 11433 | 11433 | 11433 | 11433 | 11433 | 8522 | 0
2017-07-28 19:50 | 11424 | 11424 | 11424 | 11424 | 11424 | 8522 | 0
2017-07-28 20:00 | 11410 | 11410 | 11410 | 11410 | 11410 | 8522 | 0
2017-07-28 20:10 | 11385 | 11385 | 11385 | 11385 | 11385 | 8522 | 0
2017-07-28 20:20 | 11212 | 11212 | 11212 | 11212 | 11212 | 8522 | 0
2017-07-28 20:30 | 11720 | 11720 | 11720 | 11720 | 11720 | 8522 | 0
2017-07-28 20:40 | 12085 | 12085 | 12085 | 12085 | 12085 | 8522 | 0
2017-07-28 20:50 | 12397 | 12397 | 12397 | 12397 | 12397 | 8522 | 0
2017-07-28 21:00 | 12652 | 12652 | 12652 | 12652 | 12652 | 8522 | 0
2017-07-28 21:10 | 12846 | 12846 | 12846 | 12846 | 12846 | 8522 | 0
2017-07-28 21:20 | 12982 | 12982 | 12982 | 12982 | 12982 | 8522 | 0
2017-07-28 21:30 | 13080 | 13080 | 13080 | 13080 | 13080 | 8522 | 0
2017-07-28 21:40 | 13168 | 13168 | 13168 | 13168 | 13168 | 8522 | 0
2017-07-28 21:50 | 13248 | 13248 | 13248 | 13248 | 13248 | 8523 | 0
2017-07-28 22:00 | 13317 | 13317 | 13317 | 13317 | 13317 | 8522 | 0
2017-07-28 22:10 | 13378 | 13378 | 13378 | 13378 | 13378 | 8522 | 0
2017-07-28 22:20 | 13431 | 13431 | 13431 | 13431 | 13431 | 8522 | 0
2017-07-28 22:30 | 13477 | 13477 | 13477 | 13477 | 13477 | 8523 | 0
2017-07-28 22:40 | 13529 | 13529 | 13529 | 13529 | 13529 | 8522 | 0
2017-07-28 22:50 | 13576 | 13576 | 13576 | 13576 | 13576 | 8522 | 0
2017-07-28 23:00 | 13294 | 13294 | 13294 | 13294 | 13294 | 8522 | 0
2017-07-28 23:10 | 13389 | 13389 | 13389 | 13389 | 13389 | 8522 | 0
2017-07-28 23:20 | 13113 | 13113 | 13113 | 13113 | 13113 | 8522 | 0
2017-07-28 23:30 | 13389 | 13389 | 13389 | 13389 | 13389 | 8522 | 0
2017-07-28 23:40 | 13800 | 13800 | 13800 | 13800 | 13800 | 8523 | 0
2017-07-28 23:50 | 13897 | 13897 | 13897 | 13897 | 13897 | 8523 | 0
2017-07-29 | 13903 | 13903 | 13903 | 13903 | 13903 | 8522 | 0
2017-07-29 00:10 | 13995 | 13995 | 13995 | 13995 | 13995 | 8523 | 0
2017-07-29 00:20 | 14001 | 14001 | 14001 | 14001 | 14001 | 8522 | 0
2017-07-29 00:30 | 14090 | 14090 | 14090 | 14090 | 14090 | 8523 | 0
2017-07-29 00:40 | 14266 | 14266 | 14266 | 14266 | 14266 | 8522 | 0
2017-07-29 00:50 | 14552 | 14552 | 14552 | 14552 | 14552 | 8522 | 0
2017-07-29 01:00 | 14761 | 14761 | 14761 | 14761 | 14761 | 8522 | 0
2017-07-29 01:10 | 14650 | 14650 | 14650 | 14650 | 14650 | 8522 | 0
2017-07-29 01:20 | 14571 | 14571 | 14571 | 14571 | 14571 | 8522 | 0
2017-07-29 01:30 | 14555 | 14555 | 14555 | 14555 | 14555 | 8522 | 0
2017-07-29 01:40 | 14644 | 14644 | 14644 | 14644 | 14644 | 8522 | 0
2017-07-29 01:50 | 14710 | 14710 | 14710 | 14710 | 14710 | 8522 | 0
2017-07-29 02:00 | 14694 | 14694 | 14694 | 14694 | 14694 | 8522 | 0
2017-07-29 02:10 | 14587 | 14587 | 14587 | 14587 | 14587 | 8522 | 0
2017-07-29 02:20 | 14354 | 14354 | 14354 | 14354 | 14354 | 8522 | 0
2017-07-29 02:30 | 14221 | 14221 | 14221 | 14221 | 14221 | 8522 | 0
2017-07-29 02:40 | 14145 | 14145 | 14145 | 14145 | 14145 | 8522 | 0
2017-07-29 02:50 | 14142 | 14142 | 14142 | 14142 | 14142 | 8522 | 0
2017-07-29 03:00 | 14076 | 14076 | 14076 | 14076 | 14076 | 8522 | 0
2017-07-29 03:10 | 13961 | 13961 | 13961 | 13961 | 13961 | 8522 | 0
2017-07-29 03:20 | 13852 | 13852 | 13852 | 13852 | 13852 | 8522 | 0
2017-07-29 03:30 | 13703 | 13703 | 13703 | 13703 | 13703 | 8522 | 0
2017-07-29 03:40 | 13621 | 13621 | 13621 | 13621 | 13621 | 8522 | 0
2017-07-29 03:50 | 13575 | 13575 | 13575 | 13575 | 13575 | 8522 | 0
2017-07-29 04:00 | 13849 | 13849 | 13849 | 13849 | 13849 | 8522 | 0
2017-07-29 04:10 | 14062 | 14062 | 14062 | 14062 | 14062 | 8522 | 0
2017-07-29 04:20 | 13729 | 13729 | 13729 | 13729 | 13729 | 8523 | 0
2017-07-29 04:30 | 13279 | 13279 | 13279 | 13279 | 13279 | 8522 | 0
2017-07-29 04:40 | 13365 | 13365 | 13365 | 13365 | 13365 | 8522 | 0
2017-07-29 04:50 | 13657 | 13657 | 13657 | 13657 | 13657 | 8522 | 0
2017-07-29 05:00 | 14066 | 14066 | 14066 | 14066 | 14066 | 8523 | 0
2017-07-29 05:10 | 14133 | 14133 | 14133 | 14133 | 14133 | 8522 | 0
2017-07-29 05:20 | 14226 | 14226 | 14226 | 14226 | 14226 | 8522 | 0
2017-07-29 05:30 | 14249 | 14249 | 14249 | 14249 | 14249 | 8522 | 0
2017-07-29 05:40 | 14832 | 14832 | 14832 | 14832 | 14832 | 8522 | 0
2017-07-29 05:50 | 14984 | 14984 | 14984 | 14984 | 14984 | 8522 | 0
2017-07-29 06:00 | 15421 | 15421 | 15421 | 15421 | 15421 | 8522 | 0
2017-07-29 06:10 | 15688 | 15688 | 15688 | 15688 | 15688 | 8522 | 0
2017-07-29 06:20 | 15707 | 15707 | 15707 | 15707 | 15707 | 8522 | 0
2017-07-29 06:30 | 19170 | 19170 | 19170 | 19170 | 19170 | 8522 | 0
2017-07-29 06:40 | 21036 | 21036 | 21036 | 21036 | 21036 | 8523 | 0